## Etapas
Id | Nombre | Descripción | Orden | Código
-5 | Various Protections | Various Protections | 0 | DEV
-4 | Patient Identification | Patient Identification | 1 | IDE
-3 | System Protections or Interlocks | System Protections or Interlocks | 2 | DEF
-2 | Pre-Process | Pre-Process | 3 | PRE
-1 | Quality Control | Quality Control | 4 | CON
1 | Initial Equipment Installation | Initial Equipment Installation | 5 | INS
2 | Acceptance and Commissioning | Acceptance and Commissioning | 6 | ACE
3 | Equipment Maintenance | Equipment Maintenance | 7 | MAN
4 | Treatment prescription | Treatment prescription | 8 | DEC
5 | Localization | Localization | 9 | LOC
6 | Volume Delineation | Volume Delineation | 10 | DEL
7 | Treatment Planning | Treatment Planning | 11 | PLA
8 | Post-Planning Verification | Post-Planning Verification | 12 | VER
9 | Initial Session | Initial Session | 13 | SIN
10 | Daily Treatment | Daily Treatment | 14 | TTD

## Subetapas
Id | Etapa | Nombre | Descripción | Orden
1 | 4 | Evaluation of information | Substage of 'Treatment prescription” | 0
2 | 4 | Selection of treatment goal | Substage of 'Treatment prescription” | 1
3 | 4 | Provisional prescription | Substage of 'Treatment prescription” | 2
4 | 5 | Patient identification | Substage of 'Localization” | 3
5 | 5 | Indications for imaging and immobilization proposal | Substage of 'Localization” | 4
6 | 5 | Patient preparation | Substage of 'Localization” | 5
7 | 5 | Perform immobilization | Substage of 'Localization” | 6
8 | 5 | Perform CT scan | Substage of 'Localization” | 7
9 | 5 | Image acquisition using other modalities | Substage of 'Localization” | 8
10 | 5 | Initial image review | Substage of 'Localization” | 9
11 | 5 | Document immobilization, prepare patient, technical acquisition parameters | Substage of 'Localization” | 10
12 | 5 | Image transfer | Substage of 'Localization” | 11
13 | 5 | Assign images to patient in the TPS | Substage of 'Localization” | 12
14 | 5 | Fusion and fusion validation | Substage of 'Localization” | 13
15 | 5 | Fusion validation | Substage of 'Localization” | 14
16 | 6 | Patient identification | Substage of 'Volume delineation” | 15
17 | 6 | OAR delineation | Substage of 'Volume delineation” | 16
18 | 6 | PTV delineation and OAR validation | Substage of 'Volume delineation” | 17
19 | 7 | Patient identification | Substage of 'Treatment planning” | 18
20 | 7 | Clinical prescription | Substage of 'Treatment planning” | 19
21 | 7 | Image review | Substage of 'Treatment planning” | 20
22 | 7 | Assign CT markers | Substage of 'Treatment planning” | 21
23 | 7 | Table removal/insert table model | Substage of 'Treatment Planning” | 22
24 | 7 | Density overwriting | Substage of 'Treatment Planning” | 23
25 | 7 | Selection of technical parameters | Substage of 'Treatment Planning” | 24
26 | 7 | Optimization/dose calculation | Substage of 'Treatment Planning” | 25
27 | 7 | Plan review | Substage of 'Treatment Planning” | 26
28 | 7 | Approval | Substage of 'Treatment Planning” | 27
29 | 8 | Plan recalculation on detector for treatment verification | Substage of 'Post-planning Verification” | 28
30 | 8 | Transfer of recalculation data to verification software | Substage of 'Post-planning verification” | 29
31 | 8 | Transfer of data to accelerator | Substage of 'Post-planning verification” | 30
32 | 8 | Scheduling of treatment sessions | Substage of 'Post-planning verification” | 31
33 | 8 | Treatment irradiation on detector | Substage of 'Post-planning verification” | 32
34 | 8 | Evaluation of verification results | Substage of 'Post-planning verification” | 33
35 | 8 | Acceptance of verification results | Substage of 'Post-planning verification” | 34
36 | 8 | Verification with second calculation software | Substage of 'Post-planning verification” | 35
37 | 8 | Second calculation of MU | Substage of 'Post-planning verification” | 36
38 | 8 | Planning documentation | Substage of 'Post-planning verification” | 37
39 | 8 | Review of documentation | Substage of 'Post-planning verification” | 38
40 | 8 | Treatment approval | Substage of 'Post-planning verification” | 39
41 | 9 | Preparation and review for the initial session | Substage of 'Initial Session” | 40
42 | 9 | Immobilization of the patient in the treatment position | Substage of 'Initial Session” | 41
43 | 9 | Application of displacements | Substage of 'Initial Session” | 42
44 | 9 | Acquisition of verification images | Substage of 'Initial Session” | 43
45 | 9 | Image evaluation and acceptance of displacements | Substage of 'Initial Session” | 44
46 | 9 | Application of displacements and rotations | Substage of 'Initial Session” | 45
47 | 9 | Final treatment approval | Substage of 'Initial Session” | 46
48 | 9 | Treatment | Substage of 'Initial Session” | 47
49 | 9 | Acquisition of post-treatment images | Substage of 'Initial Session” | 48
50 | 9 | Image evaluation | Substage of 'Initial Session” | 49
51 | 9 | Treatment dose verification | Substage of 'Initial Session” | 50
52 | 10 | Immobilization of the patient in the treatment position | Substage of 'Daily Treatment” | 51
53 | 10 | Application of displacements | Substage of 'Daily Treatment” | 52
54 | 10 | Acquisition of verification images | Substage of 'Daily Treatment” | 53
55 | 10 | On-site image evaluation | Substage of 'Daily Treatment” | 54
56 | 10 | Approval of displacements | Substage of 'Daily Treatment” | 55
57 | 10 | Treatment | Substage of 'Daily Treatment” | 56
58 | 10 | Treatment dose verification | Substage of 'Daily Treatment” | 57
59 | 10 | Offline image evaluation/approval | Substage of 'Daily Treatment” | 58

## Sucesos
Id | Etapa | Subetapa | Nombre | Código | Descripción | Error Humano | Frecuencia | Consecuencia | Riesgo | Completado
1 | 4 | 1 | Incorrect identification of the patient or patient data during the evaluation of information from the first consultation. | DEC-01 | Misidentifying the patient during the first consultation (e.g., assigning examination or interview data to the wrong patient, or misattributing associated reports, tests, or prior images). It is assumed that this error would affect the therapeutic intention and the patient's treatment plan. | TRUE | FB | CA | RL | Sí
2 | 4 | 3 | Omission or incorrect indication of organs at risk considered in the treatment plan on the treatment chart. | DEC-02 | Omitting or incorrectly specifying organs at risk when completing therapeutic intent data or selecting/filling a prescription template in the OIS. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
3 | 4 | 3 | Indicating a total dose prescription value, daily dose, or number of treatment sessions per week on the treatment chart that differs from the treatment plan. | DEC-03 | Indicating a total dose, daily dose, or number of sessions per week that differs from the therapeutic intent when completing the treatment plan or prescription template. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
4 | 4 | 3 | Omission or incorrect specification of dose limits for organs at risk considered in the therapeutic intent on the treatment plan. | DEC-04 | Omitting or incorrectly specifying dose limits for organs at risk when completing/selecting a prescription template in the OIS. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
5 | 4 | 3 | Omission of secondary treatment volumes included in the therapeutic intent when multiple locations are involved. | DEC-05 | Omitting secondary treatment volumes included in the clinical intent when completing the treatment plan or prescription template. It is also considered that this information is omitted from the CT chart. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FB | CA | RL | Sí
6 | 4 | 3 | Prescribing the dose for one patient to another patient. | DEC-06 | Applying the therapeutic intent of one patient to another (e.g., prescribing another patient’s dose or fractionation). It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
7 | 4 | 3 | Prescribing to an incorrect treatment volume. | DEC-07 | Selecting the wrong volume or organ for prescription (e.g., lateral error, incorrect volume, or omission/inclusion of lymph node chains). It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
8 | 4 | 3 | Failure to document special treatment considerations (e.g., pacemaker, pregnancy, prior treatments). | DEC-08 | Ignoring or omitting special treatment considerations such as pacemaker presence, pregnancy, prior treatments, concomitant therapies, comorbidities, or patient condition. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FB | CA | RL | Sí
9 | 5 | 4 | Calling the wrong patient for CT simulation. | LOC-01 | Calling the wrong patient for CT simulation. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FB | CA | RL | Sí
10 | 5 | 5 | Incomplete or incorrect instructions leading to improper immobilization or imaging of the patient. | LOC-02 | Providing incomplete or incorrect instructions that lead to improper immobilization or imaging due to protocol misapplication. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
11 | 5 | 6 | Incorrect patient preparation for imaging. | LOC-03 | Incorrect patient preparation for imaging (e.g., failure to fast for contrast, bladder filling, rectal emptying, empty stomach). It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CM | RL | Sí
12 | 5 | 6 | Performing simulation with high-contrast objects present. | LOC-04 | Performing simulation with removable high-contrast objects that degrade image quality or planning accuracy (e.g., dental prosthesis, earrings). It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CM | RL | Sí
13 | 5 | 7 | Omission, incorrect preparation, or misuse of immobilization devices during CT simulation. | LOC-05 | During CT simulation positioning, omitting or misidentifying previously prepared positioning or immobilization devices (e.g., not using an immobilizer, using an inappropriate device, mislabeling, incorrect placement, or forced positioning), or using improperly prepared devices. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
14 | 5 | 7 | Use of incorrect reference points for CT simulation. | LOC-06 | Incorrect definition of the CT origin reference (0-CT), due to omission or incorrect placement of radiopaque markers, or incorrect placement of other reference points (e.g., scars), or patient movement between positioning and CT scanning. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
15 | 5 | 7 | Incorrect patient positioning for CT simulation. | LOC-07 | uring immobilization and preparation of positioning devices on the CT simulation table, omission of positioning details or placing the patient in an uncomfortable or non-reproducible position, resulting in inaccurate CT images and inadequate positioning devices. This compromises reproducibility of treatment. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
16 | 5 | 7 | Using damaged or expired immobilization systems or immobilization products. | LOC-08 | Use of ineffective immobilization devices (e.g., looseness, poor fixation) that do not guarantee patient immobilization. It is assumed that this error would affect the development of the entire treatment of all patients. | TRUE | FM | CMA | RL | Sí
17 | 5 | 8 | Performing CT simulation with incorrect acquisition parameters. | LOC-09 | Performing CT simulation with incorrect acquisition or reconstruction parameters, due to errors in CT console data entry (e.g., scan length, field of view, slice thickness, number of slices), incorrect reconstruction settings (e.g., reconstruction thickness, metal artifact reduction, respiratory phase reconstruction), or using a non-dedicated CT scanner without an appropriate simulation protocol (e.g., differences in table design such as non-flat or sloped surfaces). It is assumed that this error would affect the entire treatment of all patients. | TRUE | FM | CA | RL | Sí
18 | 5 | 8 | CT failure resulting in erroneous images. | LOC-10 | CT malfunction producing images unsuitable for delineation or planning (e.g., excessive noise, artifacts), including undetected image degradation due to failure in quality control procedures. It is assumed that this error would affect the entire treatment of all patients. | FALSE | FB | CMA | RM | Sí
19 | 5 | 8 | Omission or incorrect acquisition fo images that take tumor motion into account. | LOC-11 | Failure to account for tumor motion during localization imaging (e.g., using rapid CT that does not capture respiratory motion, acquiring conventional CT instead of 4D CT, or incorrect acquisition due to improper placement or handling of the respiratory signal system). This error can affect volume delineation and density calculation. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FB | CA | RL | Sí
20 | 5 | 9 | Incorrect patient identification during fusion image acquisition. | LOC-12 | Entering or selecting patient identification data that does not correspond to the patient in the room, resulting in image fusion between different patients' datasets. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FB | CA | RL | Sí
21 | 5 | 9 | Incorrect patient positioning for fusion imaging (significantly different from CT setup). | LOC-13 | During acquisition of other imaging modalities (MRI, PET), positioning the patient incorrectly (e.g., arm position, immobilizers, flat board) in a way that compromises the accuracy and reliability of image fusion with CT. This error may affect volume delineation. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FB | CM | RL | Sí
22 | 5 | 9 | Incorrect placement of fiducial markers on fusion images. | LOC-14 | Incorrect placement of fiducial markers on fusion images, potentially causing errors in CT fusion and volume localization. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FB | CA | RL | Sí
23 | 5 | 9 | Performing fusion imaging with inappropriate parameters. | LOC-15 | Using inappropriate PET or MRI images due to factors such as: use of an unsuitable radiopharmaceutical for the patient’s pathology; partial volume effect caused by incorrect reconstruction filter, FOV, or matrix; use of an inappropriate MR sequence (e.g., insufficient contrast limiting lesion visualization); performing MRI with an unsuitable coil (leading to distortion and image quality loss). It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FB | CM | RL | Sí
24 | 5 | 11 | Omission or incorrect marking of reference tattoos from CT simulation on the patient or immobilization devices. | LOC-16 | Omission or incorrect tattooing of reference marks from the CT simulation on the patient’s skin, compromising reproducibility of treatment. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
25 | 5 | 11 | Omission or incorrect marking of position tattoos for respiratory tracking system components. | LOC-17 | Omission or incorrect tattooing of positions for respiratory tracking system components, preventing consistent placement throughout treatment. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FB | CA | RL | Sí
26 | 5 | 11 | Incorrect editing of CT data prior to transfer to the TPS or OIS. | LOC-18 | Incorrect data entry during CT scan editing prior to transfer to the TPS or OIS (e.g., wrong patient position such as prone/supine or head-first/feet-first). It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
27 | 5 | 11 | Omission or incorrect recording of patient positioning details from CT simulation in the treatment record. | LOC-19 | Omission or incorrect recording of patient positioning details, immobilization devices used, and other specifics from CT simulation in the treatment chart, compromising reproducibility during treatment delivery. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
28 | 5 | 12 | Failure to transfer or import images (CT, fusion, etc.) to the TPS or OIS. | LOC-20 | Failure to transfer or correctly import images (CT, fusion image, etc.) into the TPS or OIS, resulting in incomplete or altered data (e.g., incorrect number of slices, modified densities or coordinates). It is assumed that this error would affect the patient's treatment plan. | FALSE | FMB | CA | RL | Sí
29 | 5 | 13 | Assigning one patient's images to another patient in the TPS. | LOC-21 | Assigning one patient’s images to another within the TPS. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FB | CA | RL | Sí
30 | 5 | 14 | Performing image fusion with mismatched patients or different/incorrect studies. | LOC-22 | During image fusion with other modalities (MRI, PET) or previous CT scans, selecting images from the wrong patient or incorrect image sets from the same patient (e.g., wrong MRI series, previous CT). This error may affect volume delineation. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FB | CA | RL | Sí
31 | 5 | 14 | Incorrect image fusion resulting in erroneous determination of treatment volumes and organs at risk. | LOC-23 | Incorrect image fusion due to misinterpretation of anatomical morphology or use of inappropriate fusion parameters, leading to inaccurate determination of treatment volumes and organs at risk. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FB | CA | RL | Sí
32 | 6 | 17 | Omission of delineation for one, several, or all organs at risk in the TPS. | DEL-01 | Omission of delineation for one, several, or all organs at risk in the TPS during volume contouring. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
33 | 6 | 17 | Incorrect delineation or mislabeling of an organ at risk in the TPS. | DEL-02 | Incomplete delineation of an organ at risk or incorrect labeling in the TPS. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
34 | 6 | 18 | Incorrect patient identification or selection during PTV delineation. | DEL-03 | Incorrect patient identification or selection during volume delineation, resulting in contours based on the wrong patient’s clinical information. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FB | CA | RL | Sí
35 | 6 | 18 | Mislabeling of delineated treatment volumes in the TPS. | DEL-04 | Mislabeling of correctly delineated treatment volumes in the TPS, using non-standard nomenclature or ambiguous names that could lead to misinterpretation. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
36 | 6 | 18 | Omission of delineation for one or more prescribed secondary CTVs in the TPS. | DEL-05 | Omission or incomplete delineation of one or more prescribed secondary CTVs in the TPS. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
37 | 6 | 18 | Incorrect delineation of the PTV. | DEL-06 | Incorrect delineation of the PTV, such as using unvalidated automatic contouring methods, failing to incorporate tumor motion, or misassigning tumor laterality. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
38 | 7 | 19 | Incorrect patient identification when manually adding images or entering demographic data. | PLA-01 | Incorrect patient assignment during treatment planning, such as linking images or demographic data to the wrong patient. This results in using incorrect patient information for planning, which can affect the entire course of the patient's treatment. | TRUE | FB | CA | RL | Sí
39 | 7 | 20 | Incorrectly entering the total dose, dose per fraction, and number of sessions in the TPS or OIS. | PLA-02 | Errors in entering treatment parameters (total dose, dose per fraction, number of sessions) into the TPS or OIS. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
40 | 7 | 20 | Failure to update the final clinical prescription to reflect changes from the provisional prescription (e.g., total dose, fractionation) | PLA-03 | Failure to update the final clinical prescription to reflect changes from the provisional prescription, considering treatment volumes and patient immobilization. Maintaining the initial therapeutic intention without adaptation can impact the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
41 | 7 | 21 | Selection of an incorrect study for planning, such as the wrong phase of a 4D CT study. | PLA-04 | Selection of an incorrect image set for planning, such as choosing the wrong phase of a 4D CT study, incorrect phase averaging, or images acquired with a different patient position or immobilization than intended. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FB | CM | RL | Sí
42 | 7 | 22 | Incorrect assignment of the coordinate origin from CT markers or failure to assign it. | PLA-05 | Incorrect or missing assignment of the coordinate origin from CT markers, which can lead to misalignment and errors in treatment delivery. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
43 | 7 | 23 | Failure to account for treatment table absorption during planning or using a table different from the accelerator table. | PLA-06 | Failure to account for treatment table effects during planning, or incorrect consideration due to using a different table model or incorrect table positioning relative to the patient. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CM | RL | Sí
44 | 7 | 24 | Incorrect or missing correction of tissue heterogeneity densities in the treatment plan. | PLA-07 | Incorrect or missing correction of density variations in prostheses, boluses, contrast agents, imaging artifacts, etc. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CM | RL | Sí
45 | 7 | 24 | Use of an incorrect electronic density conversion curve. | PLA-08 | Use of an incorrect electron density conversion curve. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CM | RL | Sí
46 | 7 | 25 | Selection of the wrong treatment unit during the planning process. | PLA-09 | Incorrect selection of the treatment unit in the TPS, resulting in the use of data that does not correspond to the unit where the patient will be treated. Example: Using data from an accelerator in another department or from a different accelerator in the same department (different models or brands). It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
47 | 7 | 25 | Planning errors caused by inappropriate parameter settings in the treatment planning system. | PLA-10 | Incorrect planning due to inappropriate parameters, such as wrong number of arcs, erroneous scripts or templates, incorrect definition of calculation structures, inappropriate algorithm or calculation grid, incorrect dose definition and/or radiobiological objectives and/or constraints during optimization. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
48 | 7 | 25 | Incorrect definition of the patient's CT number range. | PLA-11 | Incorrect delimitation of the patient's CT number range, which may prevent recognition of immobilizers, air cavities, etc. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RM | Sí
49 | 7 | 26 | Modifying TPS usage and/or modeling procedures without validating the changes before clinical use. | PLA-12 | Use of a TPS with unvalidated or unauthorized modifications to procedures and/or modeling. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FB | CMA | RL | Sí
50 | 7 | 26 | Omission of dosimetric planning for required secondary locations. | PLA-13 | Omission of dosimetric planning for required secondary locations. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FB | CA | RL | Sí
51 | 7 | 26 | Error in considering the dose distribution of a previous phase or treatment during optimization. | PLA-14 | Incorrect consideration of dose distribution from a previous phase or treatment during optimization, leading to compensations for hot spots or cold spots of the previous phase. Each phase must be optimized independently while ensuring cumulative compliance. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
52 | 7 | 26 | TPS failure. | PLA-15 | Unexpected failure of the TPS software. It is assumed that this error would affect the entire course of the patient's treatment. | FALSE | FMB | CA | RL | Sí
53 | 7 | 27 | Failure to apply an approved change in the patient's dosimetry. | PLA-16 | Failure to perform or apply a required replanning after review or verification due to new findings that change the prescription, patient positioning or immobilization issues, anatomical changes in the patient. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
54 | 7 | 27 | Incorrect incorporation of a dosimetry change after approval. | PLA-17 | Incorrect incorporation of a dosimetry change after approval, which may occur due to new findings, patient changes, medical criteria adjustments, technical issues (e.g., collision with table or patient). It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FB | CA | RL | Sí
55 | 7 | 28 | Approval of a treatment plan different from the one jointly agreed upon by the Radiation Oncologist and the Qualified Medical Physicist. | PLA-18 | Approval of a treatment plan different from the one jointly agreed upon by the Radiation Oncologist (RTO) and the Qualified Medical Physicist (QMP). It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RM | Sí
56 | 8 | 31 | Failure to transfer electronic treatment plan information from the TPS to the OIS or to the treatment control computer | VER-01 | Failure in electronic transmission of treatment plan information from the TPS to the OIS or to the treatment control computer. Example: Incorrect transfer of multileaf collimator positions. It is assumed that this error would affect the entire course of the patient's treatment. | FALSE | FB | CA | RL | Sí
57 | 8 | 32 | Incorrect scheduling of treatment with a fractionation different from the planned regimen. | VER-02 | Incorrectly scheduling treatments that are intended to be administered every other day, weekly, etc., mistakenly scheduled for daily delivery. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
58 | 8 | 33 | Performing plan verification in treatment mode. | VER-03 | Omission of a treatment session due to incorrectly counting the verification session as a treatment session. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
59 | 8 | 34 | Incorrect modification of the initial plan based on results from an improperly performed pre-treatment verification. | VER-04 | Incorrect modification of the treatment plan due to errors in the verification process. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FB | CA | RL | Sí
60 | 9 | 41 | Omission or delay in starting treatment due to an incomplete task in the OIS in a previous stage. | SIN-01 | Omission or delay in starting a patient’s treatment due to an incomplete task in the OIS in a previous stage or a completed task that did not trigger the next step. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RM | Sí
61 | 9 | 41 | Performing treatment in a respiratory mode different from the one planned. | SIN-02 | Delivering treatment in a breathing mode different from the planned one, such as normal breathing instead of 'gating' or 'breath-hold', or vice versa. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FB | CA | RL | Sí
62 | 9 | 41 | Administering treatment using a verification plan that is not counted as a treatment session. | SIN-03 | Delivering the verification plan to the patient, resulting in an extra session that is not counted as a treatment session. This error is assumed to affect only one treatment session for the patient | TRUE | FM | CA | RL | Sí
63 | 9 | 41 | Administering treatment with a plan modified for verification purposes. | SIN-04 | Delivering treatment using a plan that is not the actual treatment plan (e.g., a plan collapsed to gantry 0°). This error is assumed to affect only one treatment session for the patient. | TRUE | FM | CA | RL | Sí
64 | 9 | 42 | Selecting the wrong patient on the treatment computer. | SIN-05 | Selecting the wrong patient during treatment preparation. It is assumed that this error would affect the entire course of the patient’s treatment. | TRUE | FM | CA | RL | Sí
65 | 9 | 42 | Incorrect patient identification at the start of treatment, or correct identification but another patient shows up by mistake. | SIN-06 | Calling the wrong patient at the start of treatment or another patient showing up by mistake. This error affects the initial treatment session and any subsequent references or adaptations. It is assumed that this error would affect the entire course of the patient’s treatment. | TRUE | FB | CA | RL | Sí
66 | 9 | 42 | Incorrect patient positioning on the treatment table (e.g., prone vs. supine, head-foot orientation error) at the start of treatment. | SIN-07 | Incorrect patient positioning on the treatment couch at the start of treatment, such as wrong orientation (prone vs. supine, head-first vs. feet-first). It is assumed that this error would affect the entire course of the patient’s treatment. | TRUE | FM | CA | RL | Sí
67 | 9 | 42 | Using immobilization devices from another patient. | SIN-08 | Incorrect patient positioning by using immobilizers from another patient. It is assumed that this error would affect the entire course of the patient’s treatment. | TRUE | FM | CA | RL | Sí
68 | 9 | 42 | Using defective patient immobilization devices. | SIN-09 | Using immobilization material that is not rigid enough, creating gaps that compromise reproducibility, including wedges, compression systems, etc, or using expired immobilization systems. This means these devices do not guarantee patient immobilization. It is assumed that this error would affect the entire course of the patient’s treatment. | TRUE | FM | CM | RL | Sí
69 | 9 | 42 | Failure to use an immobilization system or using it incorrectly. | SIN-10 | Failure to place an immobilization system or incorrect placement, such as wrong compression, incorrect indexed position, applying force to a different anatomical area. It is assumed that this error would affect the entire course of the patient’s treatment. | TRUE | FM | CA | RL | Sí
70 | 9 | 42 | Laser alignment deviation. | SIN-11 | Laser alignment deviation, causing errors in patient positioning. It is assumed that this error would affect the entire course of treatment for all patients. | FALSE | FM | CA | RL | Sí
71 | 9 | 42 | Optical distance deviation (e.g., rangefinder or reticle misalignment). | SIN-12 | Optical distance deviation (e.g., rangefinder or reticle misalignment), causing errors in patient positioning. It is assumed that this error would affect the entire course of treatment for all patients. | FALSE | FB | CA | RM | Sí
72 | 9 | 42 | Incorrect patient placement relative to the planned isocenter. | SIN-13 | Incorrect patient positioning during the initial session, such as not returning to the treatment position after image acquisition (CBCT), failing to apply or incorrectly applying necessary shifts after image acquisition. This error affects reference setup for the rest of the treatment. It is assumed that this error would affect the entire course of the patient’s treatment. | TRUE | FM | CA | RL | Sí
73 | 9 | 42 | Omission or incorrect placement of the bolus, or placement of an incorrect bolus at the start of treatment. | SIN-14 | Omission or incorrect placement of the bolus, or placement of an incorrect bolus at the first session of treatment. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FB | CA | RL | Sí
74 | 9 | 44 | Incorrect movements of mechanical elements in the IGRT system. | SIN-15 | Inaccurate movement of IGRT system mechanical elements (e.g., detector, X-ray tube), which may cause incorrect calculation of distances or magnification, misalignment between image isocenter and radiation isocenter. This error would result in incorrect patient positioning . It is assumed that this error would affect the entire course of the patient's treatment. | FALSE | FMB | CA | RL | Sí
75 | 9 | 44 | Incorrect automatic return movement of the treatment table from the verification image position to the planned isocenter. | SIN-16 | Incorrect automatic return movement of the treatment table after CBCT repositioning, especially in cases where the isocenter is highly lateralized and the system re-centers the patient for imaging. If this movement is not properly calibrated, patient positioning will be incorrect. It is assumed that this error would affect the entire course of the patient's treatment. | FALSE | FMB | CA | RL | Sí
76 | 9 | 45 | Incorrect IGRT image fusion. | SIN-17 | Incorrect merging of image fusion of IGRT images with reference images, due to misinterpretation of anatomy, inappropriate fusion parameters., etc. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
77 | 9 | 46 | Incorrect manual correction of isocenter coordinates or failure to apply necessary displacement corrections. | SIN-18 | Incorrect manual application of couch shifts after IGRT comparison with reference images, caused by misinterpretation of signs, unit conversion errors, etc. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
78 | 9 | 46 | Failure of the automatic isocenter coordinate correction system. | SIN-19 | Incorrect automatic application of couch shifts after IGRT comparison with reference images, due to equipment malfunction. It is assumed that this error would affect the entire course of the patient's treatment. | FALSE | FMB | CA | RL | Sí
79 | 9 | 46 | Incorrect manual adjustment of treatment table angle. | SIN-20 | Incorrect positioning of the treatment couch angle after locating the isocenter at the start of treatment. This error affects the initial treatment session and subsequent references/adaptations. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FB | CA | RL | Sí
80 | 9 | 47 | Incorrect final marking of the patient. | SIN-21 | Incorrect definitive marking of the patient, such as marking in the wrong location or applying incorrect shifts from CT marks. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FB | CA | RL | Sí
81 | 9 | 47 | Failure to modify the treatment plan after an approved change. | SIN-22 | Failure to incorporate approved treatment plan changes, which may result from new findings (altering PTV, dose, or fractionation), correction of previous prescription errors, equipment or technique availability issues, adjustments during initial session (e.g., collision risk), radiobiological adjustments for interruptions. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RM | Sí
82 | 9 | 47 | Incorrect incorporation of an approved treatment plan modification. | SIN-23 | After a treatment plan change has been approved (either before the first session or after several sessions), incorrectly incorporate the required modifications into the plan. Causes of the change could include new findings (which may alter PTV, dose, or fractionation), correction of previous prescription errors, lack of availability of treatment equipment or initially planned technique, changes made during the initial treatment session (small technical parameter changes such as shifts, minor variations due to collision risk, etc.), radiobiological adjustments to compensate for treatment interruptions, etc. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RM | Sí
83 | 10 | 53 | Incorrect patient identification during daily treatment, or correct identification but another patient shows up by mistake. | TTD-01 | Calling a patient other than the one scheduled for daily treatment, or a different patient showing up by mistake. This error is assumed to affect only one treatment session for one patient. | TRUE | FA | CA | RL | Sí
84 | 10 | 53 | Selecting the wrong patient on the treatment control computer. | TTD-02 | Selecting the wrong patient on the treatment control computer. This error is assumed to affect only one treatment session for one patient. | TRUE | FA | CA | RL | Sí
85 | 10 | 53 | Incorrect patient positioning on the treatment table (e.g., prone vs. supine, head-foot orientation error) during daily treatment. | TTD-03 | Incorrect immobilization of the patient during daily treatment. This error is assumed to affect only one treatment session for one patient. | TRUE | FA | CA | RL | Sí
86 | 10 | 53 | Using immobilization devices from another patient. | TTD-04 | Incorrect patient positioning on the treatment table when using another patient’s immobilizers. This error is assumed to affect only one treatment session for one patient. | TRUE | FA | CM | RL | Sí
87 | 10 | 53 | Using defective patient immobilization devices. | TTD-05 | Using ineffective immobilization devices (e.g., looseness, poor fixation) due to deterioration, meaning they do not guarantee proper immobilization. This error is assumed to affect the rest of the treatment sessions for one patient. | TRUE | FM | CA | RL | Sí
88 | 10 | 53 | Failure to place any immobilization system or incorrect placement of the system. | TTD-06 | Failure to place an immobilization system or incorrect placement, such as wrong compression, incorrect positioning in indexed compression systems, applying force to an incorrect anatomical area, not reproducing the compression force from simulation or incorrect indexing on the treatment table. It is assumed that this error would only affect one treatment session for a patient. | TRUE | FA | CM | RL | Sí
89 | 10 | 53 | Incorrect application of shifts from CT markers to the isocenter. | TTD-07 | Incorrect execution of movements to align the patient with the isocenter after immobilization and positioning based on CT marks. This error is assumed to affect only one treatment session for one patient. | TRUE | FA | CA | RL | Sí
90 | 10 | 53 | Omission or incorrect placement of boluses, or placement of incorrect boluses during daily treatment. | TTD-08 | Omission or incorrect placement of the bolus, or placement of an incorrect bolus during daily treatment. This error is assumed to affect only one treatment session for one patient. | TRUE | FB | CM | RM | Sí
91 | 10 | 53 | Laser alignment deviation. | TTD-09 | Laser alignment deviation, causing positioning error. It is assumed that this failure would affect the treatment of several patients. | FALSE | FM | CM | RL | Sí
92 | 10 | 53 | Optical distance deviation (e.g., rangefinder or reticle misalignment). | TTD-10 | Optical distance deviation (e.g., rangefinder or reticle misalignment), causing positioning errors. It is assumed that this fault would affect the treatment of several patients. | FALSE | FB | CM | RL | Sí
93 | 10 | 53 | Treating the patient with incorrect preparation. | TTD-11 | Treating the patient under conditions different from those in the simulation (e.g., dentures, bladder filling, rectal emptying). This error is assumed to affect only one treatment session for one patient. | TRUE | FM | CM | RL | Sí
94 | 10 | 55 | Incorrect positioning of breathing tracking accessories. | TTD-12 | Incorrect placement of respiratory monitoring accessories during a daily treatment session, affecting amplitude (impacting accuracy) or signal shape (if positioned more thoracic or abdominal). It is assumed that this error would affect only one treatment session for a patient. | TRUE | FM | CM | RL | Sí
95 | 10 | 55 | Incorrect movements of mechanical elements in the IGRT system. | TTD-13 | Inaccurate movement of IGRT system mechanical components (e.g., detector, X-ray tube), which may cause incorrect distance or magnification calculations, misalignment between image isocenter and radiation isocenter. It is assumed that this failure would affect the treatment of several patients. | FALSE | FB | CA | RL | Sí
96 | 10 | 55 | Incorrect automatic return movement of the treatment table from the verification image position to the planned isocenter. | TTD-14 | Incorrect automatic return movement of the treatment couch after CBCT repositioning, especially when the isocenter is highly lateralized and the system re-centers the patient for imaging. If this movement is not properly calibrated, patient positioning will be incorrect. It is assumed that this error would affect only one treatment session for a patient. | FALSE | FMB | CA | RL | Sí
97 | 10 | 56 | Incorrect execution of IGRT image fusion. | TTD-15 | Incorrect merging of IGRT images with reference images, due to misinterpretation of anatomy, inappropriate fusion parameters, etc. It is assumed that this error would only affect one treatment session for a patient. | TRUE | FM | CM | RL | Sí
98 | 10 | 56 | Incorrect manual correction of isocenter coordinates or failure to apply necessary offsets corrections. | TTD-16 | Incorrect manual application of table offsets after IGRT check caused by misinterpretation of signs, unit conversion errors, etc. This error is assumed to affect only one treatment session for one patient. | TRUE | FM | CA | RL | Sí
99 | 10 | 56 | Failure of the automatic isocenter coordinate correction system. | TTD-17 | Incorrect automatic application of couch shifts after IGRT check, due to equipment malfunction. It is assumed that this failure would affect the treatment of several patients. | FALSE | FB | CMA | RM | Sí
100 | 10 | 57 | Failure of any equipment component, leading to dosimetric and/or geometric errors during daily treatment administration. | TTD-18 | Failure of any equipment component (e.g., multileaf collimator leaves, energy, collimator, gantry, treatment couch), preventing treatment delivery with planned parameters and causing dosimetric and/or geometric errors. It is assumed that this failure would affect the treatment of several patients. | FALSE | FB | CMA | RL | Sí
101 | 10 | 57 | Administering a treatment session in a respiratory mode different from the one planned. | TTD-19 | Delivering treatment in a breathing mode different from the planned one, such as normal breathing instead of 'gating' or 'breath-hold', or vice versa. It is assumed that this error would affect only one treatment session for a patient. | TRUE | FM | CM | RL | Sí
102 | 10 | 59 | Selecting an incorrect treatment location (isocenter) on the treatment control computer. | TTD-20 | Incorrect selection of a treatment location, which occurs when a patient’s treatment includes multiple locations. It is assumed that this error would affect only one treatment session for a patient. | TRUE | FM | CA | RL | Sí
103 | 10 | 59 | Incorrect manual adjustment of treatment table angle. | TTD-21 | Incorrect manual adjustment of the treatment couch angle after locating the isocenter during daily positioning. It is assumed that this error would affect only one treatment session. | TRUE | FM | CA | RL | Sí
104 | 10 | 59 | Omission of a daily treatment session. | TTD-22 | Omission of a daily treatment session during the treatment period. It is assumed that this error would affect only one treatment session for a patient. | TRUE | FM | CM | RL | Sí
105 | 10 | 59 | Omission of an arc or field during treatment. | TTD-23 | Omission of a field or arc during treatment, for example due to irregular termination during treatment, equipment failures, patient-related issues, accidental pressing of the 'stop' key on the treatment console, etc. It is assumed that this error would affect only one treatment session for a patient. | TRUE | FA | CM | RL | Sí
106 | 10 | 59 | Incorrect completion of treatment after an interruption, due to incorrect set of the remaining MU. | TTD-24 | Failure to correctly resume treatment after an interruption, either by not administering the remaining MU or administering extra due to incorrect recording of the interruption point. It is assumed that this error would affect only one treatment session for a patient. | TRUE | FM | CM | RM | Sí
107 | 10 | 59 | Administering one or more daily treatment sessions after the prescribed number of sessions has been reached. | TTD-25 | Administering one or more additional treatment sessions after the prescribed number of sessions has been reached. It is assumed that this error would affect the entire course of the patient's treatment. | TRUE | FM | CA | RL | Sí
108 | 10 | 59 | Administering a patient’s daily treatment twice on the same day. | TTD-26 | Administering the prescribed daily treatment twice on the same day, credited in hospitalized patients. It is assumed that this error would affect only one treatment session for a patient. | TRUE | FB | CM | RM | Sí
109 | 10 | 59 | Irradiating the same arc or field twice. | TTD-27 | Irradiating the same field or arc twice in a single treatment session. It is assumed that this error would affect only one treatment session for a patient. | TRUE | FM | CM | RM | Sí
110 | 10 | 59 | ignificant (observable) patient movement, whether voluntary or involuntary. | TTD-28 | Significant (observable) patient movement, voluntary or involuntary, at any time from verification image acquisition until the end of the treatment session. It is assumed that this error would affect only one treatment session for a patient. | TRUE | FA | CM | RM | Sí

## Reductores de frecuencia
Id | Suceso | Código | Nombre | Descripción | Robustez | Marcado | Orden
1 | 52 | RF-ACE-001 | National or international protocols for the acceptance and commissioning of the TPS. | National or international protocols for the acceptance and commissioning of the TPS. | Normal | TRUE | ACE
2 | 100 | RF-ACE-002 | Protocols for adequate environmental control to prevent excessive humidity and temperature. | Protocols for environmental and utility control to ensure compliance with manufacturer specifications for accelerator operation. This includes maintaining appropriate temperature and humidity levels, as well as ensuring reliable supply of essential utilities (electricity, water) within the operating parameters recommended by the manufacturer. | Normal | TRUE | ACE
3 | 74 | RF-ACE-003 | Established protocols for initial reference states and quality controls. | Established protocols for initial reference states and quality controls. | Normal | TRUE | ACE
4 | 75 | RF-ACE-003 | Established protocols for initial reference states and quality controls. | Established protocols for initial reference states and quality controls. | Normal | TRUE | ACE
5 | 78 | RF-ACE-003 | Established protocols for initial reference states and quality controls. | Established protocols for initial reference states and quality controls. | Normal | TRUE | ACE
6 | 95 | RF-ACE-003 | Established protocols for initial reference states and quality controls. | Established protocols for initial reference states and quality controls. | Normal | TRUE | ACE
7 | 96 | RF-ACE-003 | Established protocols for initial reference states and quality controls. | Established protocols for initial reference states and quality controls. | Normal | TRUE | ACE
8 | 99 | RF-ACE-003 | Established protocols for initial reference states and quality controls. | Established protocols for initial reference states and quality controls. | Normal | TRUE | ACE
9 | 18 | RF-CON-001 | Preventive maintenance program. | Preventive maintenance program. | Normal | TRUE | CON
10 | 70 | RF-CON-001 | Preventive maintenance program. | Preventive maintenance program. | Normal | TRUE | CON
11 | 71 | RF-CON-001 | Preventive maintenance program. | Preventive maintenance program. | Normal | TRUE | CON
12 | 74 | RF-CON-001 | Preventive maintenance program. | Preventive maintenance program. | Normal | TRUE | CON
13 | 75 | RF-CON-001 | Preventive maintenance program. | Preventive maintenance program. | Normal | TRUE | CON
14 | 78 | RF-CON-001 | Preventive maintenance program. | Preventive maintenance program. | Normal | TRUE | CON
15 | 91 | RF-CON-001 | Preventive maintenance program. | Preventive maintenance program. | Normal | TRUE | CON
16 | 92 | RF-CON-001 | Preventive maintenance program. | Preventive maintenance program. | Normal | TRUE | CON
17 | 95 | RF-CON-001 | Preventive maintenance program. | Preventive maintenance program. | Normal | TRUE | CON
18 | 96 | RF-CON-001 | Preventive maintenance program. | Preventive maintenance program. | Normal | TRUE | CON
19 | 99 | RF-CON-001 | Preventive maintenance program. | Preventive maintenance program. | Normal | TRUE | CON
20 | 100 | RF-CON-001 | Preventive maintenance program. | Preventive maintenance program. | Normal | TRUE | CON
21 | 70 | RF-CON-002 | Daily quality control checks to detect parameter deviations and correct them before exceeding tolerance limits (e.g., laser alignment, MU/cGy calibration, imaging system performance). | Daily quality control checks to detect parameter deviations and correct them before exceeding tolerance limits (e.g., laser alignment, MU/cGy calibration, imaging system performance). | Normal | TRUE | CON
22 | 71 | RF-CON-002 | Daily quality control checks to detect parameter deviations and correct them before exceeding tolerance limits (e.g., laser alignment, MU/cGy calibration, imaging system performance). | Daily quality control checks to detect parameter deviations and correct them before exceeding tolerance limits (e.g., laser alignment, MU/cGy calibration, imaging system performance). | Normal | TRUE | CON
23 | 74 | RF-CON-002 | Daily quality control checks to detect parameter deviations and correct them before exceeding tolerance limits (e.g., laser alignment, MU/cGy calibration, imaging system performance). | Daily quality control checks to detect parameter deviations and correct them before exceeding tolerance limits (e.g., laser alignment, MU/cGy calibration, imaging system performance). | Normal | TRUE | CON
24 | 78 | RF-CON-002 | Daily quality control checks to detect parameter deviations and correct them before exceeding tolerance limits (e.g., laser alignment, MU/cGy calibration, imaging system performance). | Daily quality control checks to detect parameter deviations and correct them before exceeding tolerance limits (e.g., laser alignment, MU/cGy calibration, imaging system performance). | Normal | TRUE | CON
25 | 91 | RF-CON-002 | Daily quality control checks to detect parameter deviations and correct them before exceeding tolerance limits (e.g., laser alignment, MU/cGy calibration, imaging system performance). | Daily quality control checks to detect parameter deviations and correct them before exceeding tolerance limits (e.g., laser alignment, MU/cGy calibration, imaging system performance). | Normal | TRUE | CON
26 | 92 | RF-CON-002 | Daily quality control checks to detect parameter deviations and correct them before exceeding tolerance limits (e.g., laser alignment, MU/cGy calibration, imaging system performance). | Daily quality control checks to detect parameter deviations and correct them before exceeding tolerance limits (e.g., laser alignment, MU/cGy calibration, imaging system performance). | Normal | TRUE | CON
27 | 95 | RF-CON-002 | Daily quality control checks to detect parameter deviations and correct them before exceeding tolerance limits (e.g., laser alignment, MU/cGy calibration, imaging system performance). | Daily quality control checks to detect parameter deviations and correct them before exceeding tolerance limits (e.g., laser alignment, MU/cGy calibration, imaging system performance). | Normal | TRUE | CON
28 | 96 | RF-CON-002 | Daily quality control checks to detect parameter deviations and correct them before exceeding tolerance limits (e.g., laser alignment, MU/cGy calibration, imaging system performance). | Daily quality control checks to detect parameter deviations and correct them before exceeding tolerance limits (e.g., laser alignment, MU/cGy calibration, imaging system performance). | Normal | TRUE | CON
29 | 99 | RF-CON-002 | Daily quality control checks to detect parameter deviations and correct them before exceeding tolerance limits (e.g., laser alignment, MU/cGy calibration, imaging system performance). | Daily quality control checks to detect parameter deviations and correct them before exceeding tolerance limits (e.g., laser alignment, MU/cGy calibration, imaging system performance). | Normal | TRUE | CON
30 | 100 | RF-CON-002 | Daily quality control checks to detect parameter deviations and correct them before exceeding tolerance limits (e.g., laser alignment, MU/cGy calibration, imaging system performance). | Daily quality control checks to detect parameter deviations and correct them before exceeding tolerance limits (e.g., laser alignment, MU/cGy calibration, imaging system performance). | Normal | TRUE | CON
31 | 2 | RF-DEC-001 | Clinical protocols. | Written clinical protocols that formalize and standardize work procedures. | Normal | TRUE | DEC
32 | 3 | RF-DEC-001 | Clinical protocols. | Written clinical protocols that formalize and standardize work procedures. | Normal | TRUE | DEC
33 | 4 | RF-DEC-001 | Clinical protocols. | Written clinical protocols that formalize and standardize work procedures. | Normal | TRUE | DEC
34 | 5 | RF-DEC-001 | Clinical protocols. | Written clinical protocols that formalize and standardize work procedures. | Normal | TRUE | DEC
35 | 8 | RF-DEC-001 | Clinical protocols. | Written clinical protocols that formalize and standardize work procedures. | Normal | TRUE | DEC
36 | 39 | RF-DEC-001 | Clinical protocols. | Written clinical protocols that formalize and standardize work procedures. | Normal | TRUE | DEC
37 | 40 | RF-DEC-001 | Clinical protocols. | Written clinical protocols that formalize and standardize work procedures. | Normal | TRUE | DEC
38 | 57 | RF-DEC-001 | Clinical protocols. | Written clinical protocols that formalize and standardize work procedures. | Normal | TRUE | DEC
39 | 66 | RF-DEC-001 | Clinical protocols. | Written clinical protocols that formalize and standardize work procedures. | Normal | TRUE | DEC
40 | 2 | RF-DEC-002 | Automated prescription templates. | Automated prescription templates that load volumes with standardized names for delineation and predefined dose objectives. If performed manually (non-automated), robustness is reduced to Normal. | Robusta | TRUE | DEC
41 | 3 | RF-DEC-002 | Automated prescription templates. | Automated prescription templates that load volumes with standardized names for delineation and predefined dose objectives. If performed manually (non-automated), robustness is reduced to Normal. | Robusta | TRUE | DEC
42 | 4 | RF-DEC-002 | Automated prescription templates. | Automated prescription templates that load volumes with standardized names for delineation and predefined dose objectives. If performed manually (non-automated), robustness is reduced to Normal. | Robusta | TRUE | DEC
43 | 7 | RF-DEC-002 | Automated prescription templates. | Automated prescription templates that load volumes with standardized names for delineation and predefined dose objectives. If performed manually (non-automated), robustness is reduced to Normal. | Robusta | TRUE | DEC
44 | 10 | RF-DEC-002 | Automated prescription templates. | Automated prescription templates that load volumes with standardized names for delineation and predefined dose objectives. If performed manually (non-automated), robustness is reduced to Normal. | Robusta | TRUE | DEC
45 | 32 | RF-DEC-002 | Automated prescription templates. | Automated prescription templates that load volumes with standardized names for delineation and predefined dose objectives. If performed manually (non-automated), robustness is reduced to Normal. | Robusta | TRUE | DEC
46 | 33 | RF-DEC-002 | Automated prescription templates. | Automated prescription templates that load volumes with standardized names for delineation and predefined dose objectives. If performed manually (non-automated), robustness is reduced to Normal. | Robusta | TRUE | DEC
47 | 35 | RF-DEC-002 | Automated prescription templates. | Automated prescription templates that load volumes with standardized names for delineation and predefined dose objectives. If performed manually (non-automated), robustness is reduced to Normal. | Robusta | TRUE | DEC
48 | 36 | RF-DEC-002 | Automated prescription templates. | Automated prescription templates that load volumes with standardized names for delineation and predefined dose objectives. If performed manually (non-automated), robustness is reduced to Normal. | Robusta | TRUE | DEC
49 | 39 | RF-DEC-002 | Automated prescription templates. | Automated prescription templates that load volumes with standardized names for delineation and predefined dose objectives. If performed manually (non-automated), robustness is reduced to Normal. | Robusta | TRUE | DEC
50 | 1 | RF-DEC-003 | Integrated information system that allows access to all patient-related data. | Integrated system within the OIS that provides access to all patient-related information, including images and medical history. | Normal | TRUE | DEC
51 | 6 | RF-DEC-003 | Integrated information system that allows access to all patient-related data. | Integrated system within the OIS that provides access to all patient-related information, including images and medical history. | Normal | TRUE | DEC
52 | 7 | RF-DEC-003 | Integrated information system that allows access to all patient-related data. | Integrated system within the OIS that provides access to all patient-related information, including images and medical history. | Normal | TRUE | DEC
53 | 49 | RF-DEF-001 | User authentication with passwords and role-based profiles to prevent unauthorized or untrained personnel from modifying TPS parameters. | User authentication with passwords and role-based profiles to ensure only authorized and properly trained personnel can modify TPS parameters. | Normal | TRUE | DEF
54 | 102 | RF-DEL-001 | Standardized naming protocol for treatment volumes, including anatomical location in the volume name (e.g., PTV1_Dorsal, PTV2_Lumbar). | Standardized naming protocol for treatment volumes, including anatomical location in the volume name (e.g., PTV1_Dorsal, PTV2_Lumbar). | Normal | TRUE | DEL
55 | 32 | RF-DEL-002 | Automated segmentation and delineation tools. | Automated segmentation and delineation tools. | Robusta | TRUE | DEL
56 | 33 | RF-DEL-002 | Automated segmentation and delineation tools. | Automated segmentation and delineation tools. | Robusta | TRUE | DEL
57 | 35 | RF-DEL-003 | Standardized protocols or guidelines for defining treatment volumes and margins. | Standardized protocols or guidelines for defining treatment volumes and margins. | Normal | TRUE | DEL
58 | 36 | RF-DEL-003 | Standardized protocols or guidelines for defining treatment volumes and margins. | Standardized protocols or guidelines for defining treatment volumes and margins. | Normal | TRUE | DEL
59 | 37 | RF-DEL-003 | Standardized protocols or guidelines for defining treatment volumes and margins. | Standardized protocols or guidelines for defining treatment volumes and margins. | Normal | TRUE | DEL
60 | 1 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
61 | 2 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
62 | 3 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
63 | 4 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
64 | 5 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
65 | 6 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
66 | 7 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
67 | 8 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
68 | 9 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
69 | 10 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
70 | 12 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
71 | 13 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
72 | 14 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
73 | 15 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
74 | 16 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
75 | 17 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
76 | 19 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
77 | 20 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
78 | 21 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
79 | 22 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
80 | 23 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
81 | 24 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
82 | 25 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
83 | 26 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
84 | 27 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
85 | 29 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
86 | 30 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
87 | 31 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
88 | 32 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
89 | 33 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
90 | 34 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
91 | 35 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
92 | 36 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
93 | 37 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
94 | 38 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
95 | 39 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
96 | 40 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
97 | 41 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
98 | 42 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
99 | 43 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
100 | 44 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
101 | 45 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
102 | 46 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
103 | 47 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
104 | 48 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
105 | 49 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
106 | 50 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
107 | 55 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
108 | 58 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
109 | 59 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
110 | 60 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
111 | 61 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
112 | 62 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
113 | 63 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
114 | 64 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
115 | 65 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
116 | 66 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
117 | 67 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
118 | 69 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
119 | 72 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
120 | 73 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
121 | 76 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
122 | 77 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
123 | 79 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
124 | 80 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
125 | 83 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
126 | 84 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
127 | 85 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
128 | 86 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
129 | 88 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
130 | 89 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
131 | 90 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
132 | 93 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
133 | 94 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
134 | 97 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
135 | 98 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
136 | 101 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
137 | 102 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
138 | 103 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
139 | 104 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
140 | 105 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
141 | 106 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
142 | 107 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
143 | 108 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
144 | 109 | RF-DEV-001 | Maintain a moderate workload. | Maintain a balanced workload, ensuring tasks can be completed in the correct sequence and within appropriate timeframes. This includes consulting necessary information and allowing sufficient time to verify accuracy. | Normal | TRUE | DEV
145 | 1 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
146 | 2 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
147 | 3 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
148 | 4 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
149 | 5 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
150 | 6 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
151 | 7 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
152 | 8 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
153 | 9 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
154 | 10 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
155 | 12 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
156 | 13 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
157 | 14 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
158 | 15 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
159 | 16 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
160 | 17 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
161 | 19 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
162 | 20 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
163 | 21 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
164 | 22 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
165 | 23 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
166 | 24 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
167 | 25 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
168 | 26 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
169 | 27 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
170 | 29 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
171 | 30 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
172 | 31 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
173 | 32 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
174 | 33 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
175 | 34 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
176 | 35 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
177 | 36 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
178 | 37 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
179 | 38 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
180 | 39 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
181 | 40 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
182 | 41 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
183 | 42 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
184 | 43 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
185 | 44 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
186 | 45 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
187 | 46 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
188 | 47 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
189 | 48 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
190 | 49 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
191 | 50 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
192 | 55 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
193 | 58 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
194 | 59 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
195 | 60 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
196 | 61 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
197 | 62 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
198 | 63 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
199 | 64 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
200 | 65 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
201 | 66 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
202 | 67 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
203 | 69 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
204 | 72 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
205 | 73 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
206 | 76 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
207 | 77 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
208 | 79 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
209 | 80 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
210 | 83 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
211 | 84 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
212 | 85 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
213 | 86 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
214 | 88 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
215 | 89 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
216 | 90 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
217 | 93 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
218 | 94 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
219 | 97 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
220 | 98 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
221 | 101 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
222 | 102 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
223 | 103 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
224 | 104 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
225 | 105 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
226 | 106 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
227 | 107 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
228 | 108 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
229 | 109 | RF-DEV-002 | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Ensure working conditions that support concentration, such as designated areas or scheduled times without interruptions to minimize distractions. | Normal | TRUE | DEV
230 | 40 | RF-DEV-003 | Established procedure for implementing treatment changes. | Established procedure for implementing treatment changes. | Normal | TRUE | DEV
231 | 53 | RF-DEV-003 | Established procedure for implementing treatment changes. | Established procedure for implementing treatment changes. | Normal | TRUE | DEV
232 | 54 | RF-DEV-003 | Established procedure for implementing treatment changes. | Established procedure for implementing treatment changes. | Normal | TRUE | DEV
233 | 81 | RF-DEV-003 | Established procedure for implementing treatment changes. | Established procedure for implementing treatment changes. | Normal | TRUE | DEV
234 | 82 | RF-DEV-003 | Established procedure for implementing treatment changes. | Established procedure for implementing treatment changes. | Normal | TRUE | DEV
235 | 2 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
236 | 10 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
237 | 11 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
238 | 12 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
239 | 13 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
240 | 14 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
241 | 15 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
242 | 16 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
243 | 17 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
244 | 19 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
245 | 22 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
246 | 23 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
247 | 24 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
248 | 25 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
249 | 26 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
250 | 27 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
251 | 31 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
252 | 32 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
253 | 33 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
254 | 35 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
255 | 36 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
256 | 37 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
257 | 41 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
258 | 42 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
259 | 43 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
260 | 44 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
261 | 45 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
262 | 47 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
263 | 48 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
264 | 49 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
265 | 50 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
266 | 51 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
267 | 53 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
268 | 54 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
269 | 57 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
270 | 58 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
271 | 59 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
272 | 61 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
273 | 62 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
274 | 63 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
275 | 66 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
276 | 69 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
277 | 72 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
278 | 76 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
279 | 77 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
280 | 79 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
281 | 85 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
282 | 87 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
283 | 88 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
284 | 89 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
285 | 90 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
286 | 93 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
287 | 94 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
288 | 97 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
289 | 98 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
290 | 101 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
291 | 103 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
292 | 106 | RF-DEV-004 | Comprehensive training and education for all involved professionals. | Comprehensive training and education for all involved professionals. | Blanda | TRUE | DEV
293 | 105 | RF-DEV-005 | Established procedure for resuming treatment after an interruption. | Established procedure for resuming treatment after an interruption. | Normal | TRUE | DEV
294 | 106 | RF-DEV-005 | Established procedure for resuming treatment after an interruption. | Established procedure for resuming treatment after an interruption. | Normal | TRUE | DEV
295 | 1 | RF-DEV-006 | Implementation of automated workflows within the OIS. | Use of automated workflows within the OIS to link all activities associated with patient treatment. This includes access to medical records, images, and relevant clinical data. | Normal | TRUE | DEV
296 | 34 | RF-DEV-006 | Implementation of automated workflows within the OIS. | Use of automated workflows within the OIS to link all activities associated with patient treatment. This includes access to medical records, images, and relevant clinical data. | Normal | TRUE | DEV
297 | 60 | RF-DEV-006 | Implementation of automated workflows within the OIS. | Use of automated workflows within the OIS to link all activities associated with patient treatment. This includes access to medical records, images, and relevant clinical data. | Normal | TRUE | DEV
298 | 52 | RF-DEV-007 | Access to manufacturer alerts regarding TPS issues or faults. | Access to manufacturer alerts about TPS failures, allowing planning procedures to be adapted to avoid such errors. | Blanda | TRUE | DEV
299 | 65 | RF-IDE-001 | Procedure for including the patient’s photograph in the treatment chart, prepared by the RTO after the clinical prescription is finalized. | Procedure for including the patient’s photograph in the treatment chart, prepared by the RTO after the clinical prescription is finalized. | Normal | TRUE | IDE
300 | 83 | RF-IDE-001 | Procedure for including the patient’s photograph in the treatment chart, prepared by the RTO after the clinical prescription is finalized. | Procedure for including the patient’s photograph in the treatment chart, prepared by the RTO after the clinical prescription is finalized. | Normal | TRUE | IDE
301 | 13 | RF-IDE-002 | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Normal | TRUE | IDE
302 | 26 | RF-IDE-002 | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Normal | TRUE | IDE
303 | 27 | RF-IDE-002 | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Normal | TRUE | IDE
304 | 64 | RF-IDE-002 | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Normal | TRUE | IDE
305 | 65 | RF-IDE-002 | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Normal | TRUE | IDE
306 | 66 | RF-IDE-002 | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Normal | TRUE | IDE
307 | 67 | RF-IDE-002 | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Normal | TRUE | IDE
308 | 69 | RF-IDE-002 | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Normal | TRUE | IDE
309 | 73 | RF-IDE-002 | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Normal | TRUE | IDE
310 | 83 | RF-IDE-002 | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Normal | TRUE | IDE
311 | 84 | RF-IDE-002 | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Normal | TRUE | IDE
312 | 85 | RF-IDE-002 | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Normal | TRUE | IDE
313 | 86 | RF-IDE-002 | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Normal | TRUE | IDE
314 | 88 | RF-IDE-002 | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Normal | TRUE | IDE
315 | 90 | RF-IDE-002 | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Normal | TRUE | IDE
316 | 94 | RF-IDE-002 | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Procedure for inserting the patient’s positioning photograph into the electronic treatment chart. | Normal | TRUE | IDE
317 | 1 | RF-IDE-003 | Requirement for a patient photograph taken during an earlier stage. | Availability of a patient photograph taken during an earlier stage to support identification and verification throughout the treatment process. | Normal | TRUE | IDE
318 | 6 | RF-IDE-003 | Requirement for a patient photograph taken during an earlier stage. | Availability of a patient photograph taken during an earlier stage to support identification and verification throughout the treatment process. | Normal | TRUE | IDE
319 | 9 | RF-IDE-003 | Requirement for a patient photograph taken during an earlier stage. | Availability of a patient photograph taken during an earlier stage to support identification and verification throughout the treatment process. | Normal | TRUE | IDE
320 | 64 | RF-IDE-003 | Requirement for a patient photograph taken during an earlier stage. | Availability of a patient photograph taken during an earlier stage to support identification and verification throughout the treatment process. | Normal | TRUE | IDE
321 | 84 | RF-IDE-003 | Requirement for a patient photograph taken during an earlier stage. | Availability of a patient photograph taken during an earlier stage to support identification and verification throughout the treatment process. | Normal | TRUE | IDE
322 | 1 | RF-IDE-004 | Verification of patient identity using electronic identification systems. | Verification of patient identity using electronic identification systems. | Robusta | TRUE | IDE
323 | 64 | RF-IDE-004 | Verification of patient identity using electronic identification systems. | Verification of patient identity using electronic identification systems. | Robusta | TRUE | IDE
324 | 16 | RF-LOC-001 | Inspection and verification of immobilization accessories prior to use. | Inspection and verification of immobilization accessories prior to use. | Normal | TRUE | LOC
325 | 68 | RF-LOC-001 | Inspection and verification of immobilization accessories prior to use. | Inspection and verification of immobilization accessories prior to use. | Normal | TRUE | LOC
326 | 87 | RF-LOC-001 | Inspection and verification of immobilization accessories prior to use. | Inspection and verification of immobilization accessories prior to use. | Normal | TRUE | LOC
327 | 10 | RF-LOC-002 | Protocol for image acquisition and simulation documentation. | Protocol and standardized form for patient positioning, immobilization, and image acquisition for treatment planning. | Normal | TRUE | LOC
328 | 11 | RF-LOC-002 | Protocol for image acquisition and simulation documentation. | Protocol and standardized form for patient positioning, immobilization, and image acquisition for treatment planning. | Normal | TRUE | LOC
329 | 12 | RF-LOC-002 | Protocol for image acquisition and simulation documentation. | Protocol and standardized form for patient positioning, immobilization, and image acquisition for treatment planning. | Normal | TRUE | LOC
330 | 13 | RF-LOC-002 | Protocol for image acquisition and simulation documentation. | Protocol and standardized form for patient positioning, immobilization, and image acquisition for treatment planning. | Normal | TRUE | LOC
331 | 14 | RF-LOC-002 | Protocol for image acquisition and simulation documentation. | Protocol and standardized form for patient positioning, immobilization, and image acquisition for treatment planning. | Normal | TRUE | LOC
332 | 15 | RF-LOC-002 | Protocol for image acquisition and simulation documentation. | Protocol and standardized form for patient positioning, immobilization, and image acquisition for treatment planning. | Normal | TRUE | LOC
333 | 17 | RF-LOC-002 | Protocol for image acquisition and simulation documentation. | Protocol and standardized form for patient positioning, immobilization, and image acquisition for treatment planning. | Normal | TRUE | LOC
334 | 19 | RF-LOC-002 | Protocol for image acquisition and simulation documentation. | Protocol and standardized form for patient positioning, immobilization, and image acquisition for treatment planning. | Normal | TRUE | LOC
335 | 21 | RF-LOC-002 | Protocol for image acquisition and simulation documentation. | Protocol and standardized form for patient positioning, immobilization, and image acquisition for treatment planning. | Normal | TRUE | LOC
336 | 23 | RF-LOC-002 | Protocol for image acquisition and simulation documentation. | Protocol and standardized form for patient positioning, immobilization, and image acquisition for treatment planning. | Normal | TRUE | LOC
337 | 24 | RF-LOC-002 | Protocol for image acquisition and simulation documentation. | Protocol and standardized form for patient positioning, immobilization, and image acquisition for treatment planning. | Normal | TRUE | LOC
338 | 27 | RF-LOC-002 | Protocol for image acquisition and simulation documentation. | Protocol and standardized form for patient positioning, immobilization, and image acquisition for treatment planning. | Normal | TRUE | LOC
339 | 29 | RF-LOC-002 | Protocol for image acquisition and simulation documentation. | Protocol and standardized form for patient positioning, immobilization, and image acquisition for treatment planning. | Normal | TRUE | LOC
340 | 110 | RF-LOC-002 | Protocol for image acquisition and simulation documentation. | Protocol and standardized form for patient positioning, immobilization, and image acquisition for treatment planning. | Normal | TRUE | LOC
341 | 44 | RF-LOC-003 | Procedure for including warnings on the treatment chart (paper or electronic) about special circumstances such as unusual positioning, potential patient movement, presence of prostheses, contrast agents, or bolus. | Procedure for including warnings on the treatment chart (paper or electronic) about special circumstances such as unusual positioning, potential patient movement, presence of prostheses, contrast agents, or bolus. | Normal | TRUE | LOC
342 | 62 | RF-LOC-003 | Procedure for including warnings on the treatment chart (paper or electronic) about special circumstances such as unusual positioning, potential patient movement, presence of prostheses, contrast agents, or bolus. | Procedure for including warnings on the treatment chart (paper or electronic) about special circumstances such as unusual positioning, potential patient movement, presence of prostheses, contrast agents, or bolus. | Normal | TRUE | LOC
343 | 66 | RF-LOC-003 | Procedure for including warnings on the treatment chart (paper or electronic) about special circumstances such as unusual positioning, potential patient movement, presence of prostheses, contrast agents, or bolus. | Procedure for including warnings on the treatment chart (paper or electronic) about special circumstances such as unusual positioning, potential patient movement, presence of prostheses, contrast agents, or bolus. | Normal | TRUE | LOC
344 | 85 | RF-LOC-003 | Procedure for including warnings on the treatment chart (paper or electronic) about special circumstances such as unusual positioning, potential patient movement, presence of prostheses, contrast agents, or bolus. | Procedure for including warnings on the treatment chart (paper or electronic) about special circumstances such as unusual positioning, potential patient movement, presence of prostheses, contrast agents, or bolus. | Normal | TRUE | LOC
345 | 16 | RF-LOC-004 | Established procedure for purchasing and reviewing immobilization devices. | Established procedure for purchasing and reviewing immobilization devices. | Normal | TRUE | LOC
346 | 68 | RF-LOC-004 | Established procedure for purchasing and reviewing immobilization devices. | Established procedure for purchasing and reviewing immobilization devices. | Normal | TRUE | LOC
347 | 10 | RF-LOC-005 | Verification of RTT and/or nursing data prior to CT scan. | Verification of RTT and/or nursing data prior to CT scan. | Normal | TRUE | LOC
348 | 11 | RF-LOC-005 | Verification of RTT and/or nursing data prior to CT scan. | Verification of RTT and/or nursing data prior to CT scan. | Normal | TRUE | LOC
349 | 12 | RF-LOC-005 | Verification of RTT and/or nursing data prior to CT scan. | Verification of RTT and/or nursing data prior to CT scan. | Normal | TRUE | LOC
350 | 13 | RF-LOC-005 | Verification of RTT and/or nursing data prior to CT scan. | Verification of RTT and/or nursing data prior to CT scan. | Normal | TRUE | LOC
351 | 14 | RF-LOC-005 | Verification of RTT and/or nursing data prior to CT scan. | Verification of RTT and/or nursing data prior to CT scan. | Normal | TRUE | LOC
352 | 15 | RF-LOC-005 | Verification of RTT and/or nursing data prior to CT scan. | Verification of RTT and/or nursing data prior to CT scan. | Normal | TRUE | LOC
353 | 17 | RF-LOC-005 | Verification of RTT and/or nursing data prior to CT scan. | Verification of RTT and/or nursing data prior to CT scan. | Normal | TRUE | LOC
354 | 19 | RF-LOC-005 | Verification of RTT and/or nursing data prior to CT scan. | Verification of RTT and/or nursing data prior to CT scan. | Normal | TRUE | LOC
355 | 21 | RF-LOC-005 | Verification of RTT and/or nursing data prior to CT scan. | Verification of RTT and/or nursing data prior to CT scan. | Normal | TRUE | LOC
356 | 23 | RF-LOC-005 | Verification of RTT and/or nursing data prior to CT scan. | Verification of RTT and/or nursing data prior to CT scan. | Normal | TRUE | LOC
357 | 24 | RF-LOC-005 | Verification of RTT and/or nursing data prior to CT scan. | Verification of RTT and/or nursing data prior to CT scan. | Normal | TRUE | LOC
358 | 27 | RF-LOC-005 | Verification of RTT and/or nursing data prior to CT scan. | Verification of RTT and/or nursing data prior to CT scan. | Normal | TRUE | LOC
359 | 29 | RF-LOC-005 | Verification of RTT and/or nursing data prior to CT scan. | Verification of RTT and/or nursing data prior to CT scan. | Normal | TRUE | LOC
360 | 11 | RF-LOC-006 | Provide patients with clear instructions on how to prepare for imaging and/or treatment. | Provide patients with comprehensive preparation information for imaging and/or treatment, using written materials, videos, or other resources. This can be supplemented with reminders, such as phone calls, to reinforce instructions and ensure compliance. | Blanda | TRUE | LOC
361 | 21 | RF-LOC-006 | Provide patients with clear instructions on how to prepare for imaging and/or treatment. | Provide patients with comprehensive preparation information for imaging and/or treatment, using written materials, videos, or other resources. This can be supplemented with reminders, such as phone calls, to reinforce instructions and ensure compliance. | Blanda | TRUE | LOC
362 | 93 | RF-LOC-006 | Provide patients with clear instructions on how to prepare for imaging and/or treatment. | Provide patients with comprehensive preparation information for imaging and/or treatment, using written materials, videos, or other resources. This can be supplemented with reminders, such as phone calls, to reinforce instructions and ensure compliance. | Blanda | TRUE | LOC
363 | 31 | RF-LOC-007 | Implement standardized protocols for image fusion by anatomical location. | Implement standardized protocols for image fusion by anatomical location. | Normal | TRUE | LOC
364 | 76 | RF-LOC-007 | Implement standardized protocols for image fusion by anatomical location. | Implement standardized protocols for image fusion by anatomical location. | Normal | TRUE | LOC
365 | 41 | RF-LOC-008 | The RTO must clearly label the correct study for planning. | The RTO should clearly label the correct study for planning. | Normal | TRUE | LOC
366 | 30 | RF-LOC-008 | The RTO must clearly label the correct study for planning. | The RTO should clearly label the correct study for planning. | Normal | TRUE | LOC
367 | 55 | RF-LOC-008 | The RTO must clearly label the correct study for planning. | The RTO should clearly label the correct study for planning. | Normal | TRUE | LOC
368 | 25 | RF-LOC-009 | Protocol for placing respiratory control elements according to pathology. | Protocol for placing respiratory control elements according to pathology (breast, lung, etc.). | Normal | TRUE | LOC
369 | 61 | RF-LOC-009 | Protocol for placing respiratory control elements according to pathology. | Protocol for placing respiratory control elements according to pathology (breast, lung, etc.). | Normal | TRUE | LOC
370 | 94 | RF-LOC-009 | Protocol for placing respiratory control elements according to pathology. | Protocol for placing respiratory control elements according to pathology (breast, lung, etc.). | Normal | TRUE | LOC
371 | 101 | RF-LOC-009 | Protocol for placing respiratory control elements according to pathology. | Protocol for placing respiratory control elements according to pathology (breast, lung, etc.). | Normal | TRUE | LOC
372 | 94 | RF-LOC-010 | Procedure for tattooing the position of the respiratory tracking system. | Procedure for tattooing the position of the respiratory tracking system during the image acquisition stage for planning. | Normal | TRUE | LOC
373 | 46 | RF-PLA-001 | Standardized planning protocols. | Standardized planning protocols. | Normal | TRUE | PLA
374 | 47 | RF-PLA-001 | Standardized planning protocols. | Standardized planning protocols. | Normal | TRUE | PLA
375 | 50 | RF-PLA-001 | Standardized planning protocols. | Standardized planning protocols. | Normal | TRUE | PLA
376 | 51 | RF-PLA-001 | Standardized planning protocols. | Standardized planning protocols. | Normal | TRUE | PLA
377 | 55 | RF-PLA-001 | Standardized planning protocols. | Standardized planning protocols. | Normal | TRUE | PLA
378 | 41 | RF-PLA-002 | Established procedures for using the TPS. | Procedures for use (including manufacturer recommendations) for the use and modeling of TPS. This procedure may indicate the parameters and sequence of steps to be performed to plan treatment. Panama accident. | Normal | TRUE | PLA
379 | 42 | RF-PLA-002 | Established procedures for using the TPS. | Procedures for use (including manufacturer recommendations) for the use and modeling of TPS. This procedure may indicate the parameters and sequence of steps to be performed to plan treatment. Panama accident. | Normal | TRUE | PLA
380 | 43 | RF-PLA-002 | Established procedures for using the TPS. | Procedures for use (including manufacturer recommendations) for the use and modeling of TPS. This procedure may indicate the parameters and sequence of steps to be performed to plan treatment. Panama accident. | Normal | TRUE | PLA
381 | 44 | RF-PLA-002 | Established procedures for using the TPS. | Procedures for use (including manufacturer recommendations) for the use and modeling of TPS. This procedure may indicate the parameters and sequence of steps to be performed to plan treatment. Panama accident. | Normal | TRUE | PLA
382 | 45 | RF-PLA-002 | Established procedures for using the TPS. | Procedures for use (including manufacturer recommendations) for the use and modeling of TPS. This procedure may indicate the parameters and sequence of steps to be performed to plan treatment. Panama accident. | Normal | TRUE | PLA
383 | 48 | RF-PLA-002 | Established procedures for using the TPS. | Procedures for use (including manufacturer recommendations) for the use and modeling of TPS. This procedure may indicate the parameters and sequence of steps to be performed to plan treatment. Panama accident. | Normal | TRUE | PLA
384 | 49 | RF-PLA-002 | Established procedures for using the TPS. | Procedures for use (including manufacturer recommendations) for the use and modeling of TPS. This procedure may indicate the parameters and sequence of steps to be performed to plan treatment. Panama accident. | Normal | TRUE | PLA
385 | 51 | RF-PLA-002 | Established procedures for using the TPS. | Procedures for TPS use and modeling, incorporating manufacturer recommendations. These should define parameters and the sequence of steps required for accurate treatment planning. (Reference: Panama accident) | Normal | TRUE | PLA
386 | 49 | RF-PLA-003 | Explicit prohibition on making changes to TPS usage procedures if they are not validated. | Explicit prohibition of unauthorized TPS usage, including operating in unapproved modes or performing actions not supported by the manufacturer (e.g., modifying files, changing file system locations). (Reference: Panama accident) | Normal | TRUE | PLA
387 | 43 | RF-PLA-004 | Automatic assignment of the treatment table with correct density values in the TPS. | Automatic assignment of the treatment table with correct density values in the TPS. | Robusta | TRUE | PLA
388 | 45 | RF-PLA-005 | Automatic identification of the image source and linking to the corresponding electron density conversion curve. | Automatic identification of the image source and linking to the corresponding electron density conversion curve. | Robusta | TRUE | PLA
389 | 39 | RF-PLA-006 | Procedure for using automatic planning templates. | Procedure for utilizing planning templates that automatically apply dose to the PTV and organs at risk. This process may also include the creation of planning volumes, tables, calculation grids, and beam arrangements. | Muy robusta | TRUE | PLA
390 | 43 | RF-PLA-006 | Procedure for using automatic planning templates. | Procedure for utilizing planning templates that automatically apply dose to the PTV and organs at risk. This process may also include the creation of planning volumes, tables, calculation grids, and beam arrangements. | Muy robusta | TRUE | PLA
391 | 47 | RF-PLA-006 | Procedure for using automatic planning templates. | Procedure for utilizing planning templates that automatically apply dose to the PTV and organs at risk. This process may also include the creation of planning volumes, tables, calculation grids, and beam arrangements. | Muy robusta | TRUE | PLA
392 | 89 | RF-SIN-001 | Use of automatic table movement. | Use of automatic table movement. | Muy robusta | TRUE | SIN
393 | 80 | RF-SIN-002 | Procedure establishing that the RTT who marks the patient's coordinates with a marker is the same person who tattoos them. | Procedure establishing that the RTT who marks the patient's coordinates with a marker is the same person who tattoos them. | Normal | TRUE | SIN
394 | 105 | RF-SIN-003 | Procedure for defining a logical sequence of treatment fields within the accelerator. | Procedure for defining a logical sequence of treatment fields within the accelerator. | Normal | TRUE | SIN
395 | 109 | RF-SIN-003 | Procedure for defining a logical sequence of treatment fields within the accelerator. | Procedure for defining a logical sequence of treatment fields within the accelerator. | Normal | TRUE | SIN
396 | 105 | RF-SIN-004 | Automatic irradiation sequence for treatment fields. | Automatic sequencing of fields to process all of them collectively, eliminating the need for individual loading. | Normal | TRUE | SIN
397 | 109 | RF-SIN-004 | Automatic irradiation sequence for treatment fields. | Automatic sequencing of fields to process all of them collectively, eliminating the need for individual loading. | Normal | TRUE | SIN
398 | 76 | RF-SIN-005 | IGRT protocols organized by pathology. | Protocol specifying appropriate merge parameters, methodology, and verification options to ensure correct merging. | Normal | TRUE | SIN
399 | 97 | RF-SIN-005 | IGRT protocols organized by pathology. | Protocol specifying appropriate merge parameters, methodology, and verification options to ensure correct merging. | Normal | TRUE | SIN
400 | 73 | RF-SIN-006 | Use of a bolus field. | Use of a field shaped like the bolus on the patient, generated from the planning system. The light simulation guides accurate positioning of the treatment bolus. | Normal | TRUE | SIN
401 | 90 | RF-SIN-006 | Use of a bolus field. | Use of a field shaped like the bolus on the patient, generated from the planning system. The light simulation guides accurate positioning of the treatment bolus. | Normal | TRUE | SIN
402 | 25 | RF-TTD-001 | Procedure ensuring the continuous presence of two RTTs at the machine to enable redundant monitoring of the treatment process. | Procedure ensuring the continuous presence of two RTTs at the machine, enabling redundant monitoring of treatment parameters, CCD cameras, and intercom systems | Normal | TRUE | TTD
403 | 57 | RF-TTD-001 | Procedure ensuring the continuous presence of two RTTs at the machine to enable redundant monitoring of the treatment process. | Procedure ensuring the continuous presence of two RTTs at the machine, enabling redundant monitoring of treatment parameters, CCD cameras, and intercom systems | Normal | TRUE | TTD
404 | 61 | RF-TTD-001 | Procedure ensuring the continuous presence of two RTTs at the machine to enable redundant monitoring of the treatment process. | Procedure ensuring the continuous presence of two RTTs at the machine, enabling redundant monitoring of treatment parameters, CCD cameras, and intercom systems | Normal | TRUE | TTD
405 | 62 | RF-TTD-001 | Procedure ensuring the continuous presence of two RTTs at the machine to enable redundant monitoring of the treatment process. | Procedure ensuring the continuous presence of two RTTs at the machine, enabling redundant monitoring of treatment parameters, CCD cameras, and intercom systems | Normal | TRUE | TTD
406 | 63 | RF-TTD-001 | Procedure ensuring the continuous presence of two RTTs at the machine to enable redundant monitoring of the treatment process. | Procedure ensuring the continuous presence of two RTTs at the machine, enabling redundant monitoring of treatment parameters, CCD cameras, and intercom systems | Normal | TRUE | TTD
407 | 77 | RF-TTD-001 | Procedure ensuring the continuous presence of two RTTs at the machine to enable redundant monitoring of the treatment process. | Procedure ensuring the continuous presence of two RTTs at the machine, enabling redundant monitoring of treatment parameters, CCD cameras, and intercom systems | Normal | TRUE | TTD
408 | 94 | RF-TTD-001 | Procedure ensuring the continuous presence of two RTTs at the machine to enable redundant monitoring of the treatment process. | Procedure ensuring the continuous presence of two RTTs at the machine, enabling redundant monitoring of treatment parameters, CCD cameras, and intercom systems | Normal | TRUE | TTD
409 | 98 | RF-TTD-001 | Procedure ensuring the continuous presence of two RTTs at the machine to enable redundant monitoring of the treatment process. | Procedure ensuring the continuous presence of two RTTs at the machine, enabling redundant monitoring of treatment parameters, CCD cameras, and intercom systems | Normal | TRUE | TTD
410 | 101 | RF-TTD-001 | Procedure ensuring the continuous presence of two RTTs at the machine to enable redundant monitoring of the treatment process. | Procedure ensuring the continuous presence of two RTTs at the machine, enabling redundant monitoring of treatment parameters, CCD cameras, and intercom systems | Normal | TRUE | TTD
411 | 105 | RF-TTD-001 | Procedure ensuring the continuous presence of two RTTs at the machine to enable redundant monitoring of the treatment process. | Procedure ensuring the continuous presence of two RTTs at the machine, enabling redundant monitoring of treatment parameters, CCD cameras, and intercom systems | Normal | TRUE | TTD
412 | 108 | RF-TTD-001 | Procedure ensuring the continuous presence of two RTTs at the machine to enable redundant monitoring of the treatment process. | Procedure ensuring the continuous presence of two RTTs at the machine, enabling redundant monitoring of treatment parameters, CCD cameras, and intercom systems | Normal | TRUE | TTD
413 | 109 | RF-TTD-001 | Procedure ensuring the continuous presence of two RTTs at the machine to enable redundant monitoring of the treatment process. | Procedure ensuring the continuous presence of two RTTs at the machine, enabling redundant monitoring of treatment parameters, CCD cameras, and intercom systems | Normal | TRUE | TTD
414 | 110 | RF-TTD-001 | Procedure ensuring the continuous presence of two RTTs at the machine to enable redundant monitoring of the treatment process. | Procedure ensuring the continuous presence of two RTTs at the machine, enabling redundant monitoring of treatment parameters, CCD cameras, and intercom systems | Normal | TRUE | TTD
415 | 90 | RF-TTD-002 | Procedure for establishing a system to identify the bolus using the patient’s data. | Procedure for establishing a system to identify the bolus using the patient’s data. | Normal | TRUE | TTD
416 | 73 | RF-TTD-003 | Procedure ensuring that the presence of a bolus is clearly indicated by a note in the corresponding field, so the information is visible when downloaded via the network. | Procedure ensuring that the presence of a bolus is clearly indicated by a note in the corresponding field, making it visible when downloaded via the network. | Normal | TRUE | TTD
417 | 90 | RF-TTD-003 | Procedure ensuring that the presence of a bolus is clearly indicated by a note in the corresponding field, so the information is visible when downloaded via the network. | Procedure ensuring that the presence of a bolus is clearly indicated by a note in the corresponding field, making it visible when downloaded via the network. | Normal | TRUE | TTD
418 | 57 | RF-TTD-004 | Procedure requiring the RTT to systematically review the treatment schedule. | Procedure requiring the RTT to systematically review the treatment schedule. | Normal | TRUE | TTD
419 | 104 | RF-TTD-004 | Procedure requiring the RTT to systematically review the treatment schedule. | Procedure requiring the RTT to systematically review the treatment schedule. | Normal | TRUE | TTD
420 | 107 | RF-TTD-004 | Procedure requiring the RTT to systematically review the treatment schedule. | Procedure requiring the RTT to systematically review the treatment schedule. | Normal | TRUE | TTD
421 | 108 | RF-TTD-004 | Procedure requiring the RTT to systematically review the treatment schedule. | Procedure requiring the RTT to systematically review the treatment schedule. | Normal | TRUE | TTD
422 | 102 | RF-TTD-005 | Procedure establishing a consistent order of irradiation for treatment locations during each session (e.g., always treating the volume closest to the head first, followed by the one closest to the feet). | Procedure establishing a consistent order of irradiation for treatment locations during each session (e.g., always treating the volume closest to the head first, followed by the one closest to the feet). | Normal | TRUE | TTD
423 | 110 | RF-TTD-006 | Instructions for the patient to remain completely still during treatment. | Instructions for the patient to remain completely still during treatment. | Blanda | TRUE | TTD
424 | 59 | RF-VER-001 | Procedure prohibiting manual modification of treatment plans on the OIS. | Procedure prohibiting manual modification of treatment plans on the OIS. | Normal | TRUE | VER
425 | 58 | RF-VER-002 | Protocol requiring that, once verification of the plan is complete, the verification plan is either deauthorized or deleted. | Protocol requiring that, once verification of the plan is complete, the verification plan is either deauthorized or deleted. | Normal | TRUE | VER
426 | 62 | RF-VER-002 | Protocol requiring that, once verification of the plan is complete, the verification plan is either deauthorized or deleted. | Protocol requiring that, once verification of the plan is complete, the verification plan is either deauthorized or deleted. | Normal | TRUE | VER
427 | 63 | RF-VER-002 | Protocol requiring that, once verification of the plan is complete, the verification plan is either deauthorized or deleted. | Protocol requiring that, once verification of the plan is complete, the verification plan is either deauthorized or deleted. | Normal | TRUE | VER
428 | 58 | RF-VER-003 | Procedure ensuring that the RTO does not approve treatment until the plan has been fully verified. | Procedure ensuring that the RTO does not approve treatment until the plan has been fully verified. Consequently, the verification plan cannot be delivered in treatment mode. | Normal | TRUE | VER

## Barreras
Id | Suceso | Código | Nombre | Descripción | Robustez | Marcado | Orden
429 | 1 | BA-DEC-001 | Internal Peer Review of RTOs. | Presentation and discussion of new patients during RTO clinical sessions. | Normal | TRUE | DEC
430 | 6 | BA-DEC-001 | Internal Peer Review of RTOs. | Presentation and discussion of new patients during RTO clinical sessions. | Normal | TRUE | DEC
431 | 7 | BA-DEC-001 | Internal Peer Review of RTOs. | Presentation and discussion of new patients during RTO clinical sessions. | Normal | TRUE | DEC
432 | 8 | BA-DEC-001 | Internal Peer Review of RTOs. | Presentation and discussion of new patients during RTO clinical sessions. | Normal | TRUE | DEC
433 | 40 | BA-DEC-001 | Internal Peer Review of RTOs. | Presentation and discussion of new patients during RTO clinical sessions. | Normal | TRUE | DEC
434 | 2 | BA-DEC-002 | Automatic comparison of the data entered with the planning protocol. | The planning protocol requires specific ROIs. Automatic comparison (via scripts/templates) of the data entered for planning against the protocol helps detect prescription errors, such as incorrect doses or misnamed PTVs and organs at risk. | Robusta | TRUE | DEC
435 | 3 | BA-DEC-002 | Automatic comparison of the data entered with the planning protocol. | The planning protocol requires specific ROIs. Automatic comparison (via scripts/templates) of the data entered for planning against the protocol helps detect prescription errors, such as incorrect doses or misnamed PTVs and organs at risk. | Robusta | TRUE | DEC
436 | 4 | BA-DEC-002 | Automatic comparison of the data entered with the planning protocol. | The planning protocol requires specific ROIs. Automatic comparison (via scripts/templates) of the data entered for planning against the protocol helps detect prescription errors, such as incorrect doses or misnamed PTVs and organs at risk. | Robusta | TRUE | DEC
437 | 5 | BA-DEC-002 | Automatic comparison of the data entered with the planning protocol. | The planning protocol requires specific ROIs. Automatic comparison (via scripts/templates) of the data entered for planning against the protocol helps detect prescription errors, such as incorrect doses or misnamed PTVs and organs at risk. | Robusta | TRUE | DEC
438 | 7 | BA-DEC-002 | Automatic comparison of the data entered with the planning protocol. | The planning protocol requires specific ROIs. Automatic comparison (via scripts/templates) of the data entered for planning against the protocol helps detect prescription errors, such as incorrect doses or misnamed PTVs and organs at risk. | Robusta | TRUE | DEC
439 | 32 | BA-DEC-002 | Automatic comparison of the data entered with the planning protocol. | The planning protocol requires specific ROIs. Automatic comparison (via scripts/templates) of the data entered for planning against the protocol helps detect prescription errors, such as incorrect doses or misnamed PTVs and organs at risk. | Robusta | TRUE | DEC
440 | 35 | BA-DEC-002 | Automatic comparison of the data entered with the planning protocol. | The planning protocol requires specific ROIs. Automatic comparison (via scripts/templates) of the data entered for planning against the protocol helps detect prescription errors, such as incorrect doses or misnamed PTVs and organs at risk. | Robusta | TRUE | DEC
441 | 40 | BA-DEC-002 | Automatic comparison of the data entered with the planning protocol. | The planning protocol requires specific ROIs. Automatic comparison (via scripts/templates) of the data entered for planning against the protocol helps detect prescription errors, such as incorrect doses or misnamed PTVs and organs at risk. | Robusta | TRUE | DEC
442 | 1 | BA-DEC-003 | Automatic retrieval of the patient’s digitized medical history. | Automatic retrieval of the patient’s digitized medical history by reading their health card, ID card, or similar identifier. | Muy robusta | TRUE | DEC
443 | 74 | BA-DEC-003 | Automatic retrieval of the patient’s digitized medical history. | Automatic retrieval of the patient’s digitized medical history by reading their health card, ID card, or similar identifier. | Muy robusta | TRUE | DEC
444 | 95 | BA-DEC-003 | Automatic retrieval of the patient’s digitized medical history. | Automatic retrieval of the patient’s digitized medical history by reading their health card, ID card, or similar identifier. | Muy robusta | TRUE | DEC
445 | 107 | BA-DEF-001 | Interlock of the R&V system. | Interlocking of the R&V system. This electronic system records and verifies all information about the treatment administered to a patient, compares it with the planned and scheduled treatment, and prevents treatment if inconsistencies are detected. | Muy robusta | TRUE | DEF
446 | 108 | BA-DEF-001 | Interlock of the R&V system. | Interlocking of the R&V system. This electronic system records and verifies all information about the treatment administered to a patient, compares it with the planned and scheduled treatment, and prevents treatment if inconsistencies are detected. | Muy robusta | TRUE | DEF
447 | 109 | BA-DEF-001 | Interlock of the R&V system. | Interlocking of the R&V system. This electronic system records and verifies all information about the treatment administered to a patient, compares it with the planned and scheduled treatment, and prevents treatment if inconsistencies are detected. | Muy robusta | TRUE | DEF
448 | 104 | BA-DEF-002 | Notification generated by the R&V system. | Notification or alarm from the recording and verification system. Electronic system that records and verifies all information about the treatment administered to a patient and displays a notification or alarm if it detects inconsistencies. | Robusta | TRUE | DEF
449 | 105 | BA-DEF-002 | Notification generated by the R&V system. | Warning or alarm generated by the R&V system. This electronic system records and verifies all information about the treatment administered to a patient and issues alerts if inconsistencies are detected. | Robusta | TRUE | DEF
450 | 106 | BA-DEF-002 | Notification generated by the R&V system. | Warning or alarm generated by the R&V system. This electronic system records and verifies all information about the treatment administered to a patient and issues alerts if inconsistencies are detected. | Robusta | TRUE | DEF
451 | 79 | BA-DEF-003 | Interlock for incorrect treatment table angle. | Interlock mechanism that prevents irradiation if the treatment table angle does not match the angle defined in the treatment plan. | Muy robusta | TRUE | DEF
452 | 103 | BA-DEF-003 | Interlock for incorrect treatment table angle. | Interlock that prevents irradiation if the treatment table position does not match the position defined in the treatment plan or initial session. This barrier applies only when the table position has been previously defined. | Muy robusta | TRUE | DEF
453 | 64 | BA-DEF-004 | Interlock due to an error in the treatment table position. | Interlock that prevents irradiation if the treatment table position does not match the position defined in the treatment plan or initial session. This barrier applies only when the table position has been previously defined. | Muy robusta | TRUE | DEF
454 | 65 | BA-DEF-004 | Interlock due to an error in the treatment table position. | Interlock that prevents irradiation if the treatment table position does not match the position defined in the treatment plan or initial session. This barrier applies only when the table position has been previously defined. | Muy robusta | TRUE | DEF
455 | 66 | BA-DEF-004 | Interlock due to an error in the treatment table position. | Interlock that prevents irradiation if the treatment table position does not match the position defined in the treatment plan or initial session. This barrier applies only when the table position has been previously defined. | Muy robusta | TRUE | DEF
456 | 69 | BA-DEF-004 | Interlock due to an error in the treatment table position. | Interlock that prevents irradiation if the treatment table position does not match the position defined in the treatment plan or initial session. This barrier applies only when the table position has been previously defined. | Muy robusta | TRUE | DEF
457 | 70 | BA-DEF-004 | Interlock due to an error in the treatment table position. | Interlock that prevents irradiation if the treatment table position does not match the position defined in the treatment plan or initial session. This barrier applies only when the table position has been previously defined. | Muy robusta | TRUE | DEF
458 | 72 | BA-DEF-004 | Interlock due to an error in the treatment table position. | Interlock that prevents irradiation if the treatment table position does not match the position defined in the treatment plan or initial session. This barrier applies only when the table position has been previously defined. | Muy robusta | TRUE | DEF
459 | 75 | BA-DEF-004 | Interlock due to an error in the treatment table position. | Interlock that prevents irradiation if the treatment table position does not match the position defined in the treatment plan or initial session. This barrier applies only when the table position has been previously defined. | Muy robusta | TRUE | DEF
460 | 77 | BA-DEF-004 | Interlock due to an error in the treatment table position. | Interlock that prevents irradiation if the treatment table position does not match the position defined in the treatment plan or initial session. This barrier applies only when the table position has been previously defined. | Muy robusta | TRUE | DEF
461 | 78 | BA-DEF-004 | Interlock due to an error in the treatment table position. | Interlock that prevents irradiation if the treatment table position does not match the position defined in the treatment plan or initial session. This barrier applies only when the table position has been previously defined. | Muy robusta | TRUE | DEF
462 | 83 | BA-DEF-004 | Interlock due to an error in the treatment table position. | Interlock that prevents irradiation if the treatment table position does not match the position defined in the treatment plan or initial session. This barrier applies only when the table position has been previously defined. | Muy robusta | TRUE | DEF
463 | 84 | BA-DEF-004 | Interlock due to an error in the treatment table position. | Interlock that prevents irradiation if the treatment table position does not match the position defined in the treatment plan or initial session. This barrier applies only when the table position has been previously defined. | Muy robusta | TRUE | DEF
464 | 85 | BA-DEF-004 | Interlock due to an error in the treatment table position. | Interlock that prevents irradiation if the treatment table position does not match the position defined in the treatment plan or initial session. This barrier applies only when the table position has been previously defined. | Muy robusta | TRUE | DEF
465 | 88 | BA-DEF-004 | Interlock due to an error in the treatment table position. | Interlock that prevents irradiation if the treatment table position does not match the position defined in the treatment plan or initial session. This barrier applies only when the table position has been previously defined. | Muy robusta | TRUE | DEF
466 | 89 | BA-DEF-004 | Interlock due to an error in the treatment table position. | Interlock that prevents irradiation if the treatment table position does not match the position defined in the treatment plan or initial session. This barrier applies only when the table position has been previously defined. | Muy robusta | TRUE | DEF
467 | 96 | BA-DEF-004 | Interlock due to an error in the treatment table position. | Interlock that prevents irradiation if the treatment table position does not match the position defined in the treatment plan or initial session. This barrier applies only when the table position has been previously defined. | Muy robusta | TRUE | DEF
468 | 98 | BA-DEF-004 | Interlock due to an error in the treatment table position. | Interlock that prevents irradiation if the treatment table position does not match the position defined in the treatment plan or initial session. This barrier applies only when the table position has been previously defined. | Muy robusta | TRUE | DEF
469 | 99 | BA-DEF-004 | Interlock due to an error in the treatment table position. | Interlock that prevents irradiation if the treatment table position does not match the position defined in the treatment plan or initial session. This barrier applies only when the table position has been previously defined. | Muy robusta | TRUE | DEF
470 | 102 | BA-DEF-004 | Interlock due to an error in the treatment table position. | Interlock that prevents irradiation if the treatment table position does not match the position defined in the treatment plan or initial session. This barrier applies only when the table position has been previously defined. | Muy robusta | TRUE | DEF
471 | 28 | BA-DEF-005 | Software alert indicating incorrect transfer or import of data or images (e.g., from CT to TPS, or from TPS to the OIS). | Software alert indicating incorrect transfer or import of data or images (e.g., from CT to TPS or from TPS to the OIS). If the system not only alerts but also rejects erroneous data, this barrier must be VERY ROBUST. | Robusta | TRUE | DEF
472 | 56 | BA-DEF-005 | Software alert indicating incorrect transfer or import of data or images (e.g., from CT to TPS, or from TPS to the OIS). | Software alert indicating incorrect transfer or import of data or images (e.g., from CT to TPS or from TPS to the OIS). If the system not only alerts but also rejects erroneous data, this barrier must be VERY ROBUST. | Robusta | TRUE | DEF
473 | 60 | BA-DEF-006 | OIS alert sent to the person responsible for the task. | OIS alert sent to the task manager and/or patient flow manager when the time requiered to perform the task exceeds the scheduled time. | Robusta | TRUE | DEF
474 | 74 | BA-DEF-007 | Imaging system interlock. | Interlock of the imaging system that prevents its use if the image receptor and X-ray tube are not correctly positioned. | Muy robusta | TRUE | DEF
475 | 95 | BA-DEF-007 | Imaging system interlock. | Interlock of the imaging system that prevents its use if the image receptor and X-ray tube are not correctly positioned. | Muy robusta | TRUE | DEF
476 | 74 | BA-DEF-008 | Equipment interlock that prevents beam activation or stops irradiation in the event of a component failure causing dosimetric or geometric parameters to deviate beyond tolerance limits. | Equipment interlock that prevents beam activation or stops irradiation in the event of a component failure causing dosimetric or geometric parameters to deviate beyond tolerance limits. Examples include interlocks for gantry position errors, beam deflection system failures, or MLC position errors. | Muy robusta | TRUE | DEF
477 | 95 | BA-DEF-008 | Equipment interlock that prevents beam activation or stops irradiation in the event of a component failure causing dosimetric or geometric parameters to deviate beyond tolerance limits. | Equipment interlock that prevents beam activation or stops irradiation in the event of a component failure causing dosimetric or geometric parameters to deviate beyond tolerance limits. Examples include interlocks for gantry position errors, beam deflection system failures, or MLC position errors. | Muy robusta | TRUE | DEF
478 | 100 | BA-DEF-008 | Equipment interlock that prevents beam activation or stops irradiation in the event of a component failure causing dosimetric or geometric parameters to deviate beyond tolerance limits. | Equipment interlock that prevents beam activation or stops irradiation in the event of a component failure causing dosimetric or geometric parameters to deviate beyond tolerance limits. Examples include interlocks for gantry position errors, beam deflection system failures, or MLC position errors. | Muy robusta | TRUE | DEF
479 | 100 | BA-DEF-009 | Equipment interlock that prevents beam activation or stops irradiation in the event of a component failure causing dosimetric or geometric parameters to deviate beyond tolerance limits. | Equipment interlock that prevents beam activation or stops irradiation in the event of a component failure causing dosimetric or geometric parameters to deviate beyond tolerance limits. Examples include interlocks for gantry position errors, beam deflection system failures, or MLC position errors. | Muy robusta | TRUE | DEF
480 | 2 | BA-DEL-001 | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. This step helps detect errors from previous stages, such as treatment prescription or anatomical data collection, and may include checks for adequate slice thickness and absence of artifacts. | Blanda | TRUE | DEL
481 | 4 | BA-DEL-001 | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. This step helps detect errors from previous stages, such as treatment prescription or anatomical data collection, and may include checks for adequate slice thickness and absence of artifacts. | Blanda | TRUE | DEL
482 | 7 | BA-DEL-001 | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. This step helps detect errors from previous stages, such as treatment prescription or anatomical data collection, and may include checks for adequate slice thickness and absence of artifacts. | Blanda | TRUE | DEL
483 | 9 | BA-DEL-001 | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. This step helps detect errors from previous stages, such as treatment prescription or anatomical data collection, and may include checks for adequate slice thickness and absence of artifacts. | Blanda | TRUE | DEL
484 | 10 | BA-DEL-001 | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. This step helps detect errors from previous stages, such as treatment prescription or anatomical data collection, and may include checks for adequate slice thickness and absence of artifacts. | Blanda | TRUE | DEL
485 | 11 | BA-DEL-001 | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. This step helps detect errors from previous stages, such as treatment prescription or anatomical data collection, and may include checks for adequate slice thickness and absence of artifacts. | Blanda | TRUE | DEL
486 | 12 | BA-DEL-001 | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. This step helps detect errors from previous stages, such as treatment prescription or anatomical data collection, and may include checks for adequate slice thickness and absence of artifacts. | Blanda | TRUE | DEL
487 | 15 | BA-DEL-001 | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. This step helps detect errors from previous stages, such as treatment prescription or anatomical data collection, and may include checks for adequate slice thickness and absence of artifacts. | Blanda | TRUE | DEL
488 | 17 | BA-DEL-001 | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. This step helps detect errors from previous stages, such as treatment prescription or anatomical data collection, and may include checks for adequate slice thickness and absence of artifacts. | Blanda | TRUE | DEL
489 | 18 | BA-DEL-001 | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. This step helps detect errors from previous stages, such as treatment prescription or anatomical data collection, and may include checks for adequate slice thickness and absence of artifacts. | Blanda | TRUE | DEL
490 | 20 | BA-DEL-001 | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. This step helps detect errors from previous stages, such as treatment prescription or anatomical data collection, and may include checks for adequate slice thickness and absence of artifacts. | Blanda | TRUE | DEL
491 | 21 | BA-DEL-001 | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. This step helps detect errors from previous stages, such as treatment prescription or anatomical data collection, and may include checks for adequate slice thickness and absence of artifacts. | Blanda | TRUE | DEL
492 | 22 | BA-DEL-001 | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. This step helps detect errors from previous stages, such as treatment prescription or anatomical data collection, and may include checks for adequate slice thickness and absence of artifacts. | Blanda | TRUE | DEL
493 | 23 | BA-DEL-001 | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. This step helps detect errors from previous stages, such as treatment prescription or anatomical data collection, and may include checks for adequate slice thickness and absence of artifacts. | Blanda | TRUE | DEL
494 | 26 | BA-DEL-001 | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. This step helps detect errors from previous stages, such as treatment prescription or anatomical data collection, and may include checks for adequate slice thickness and absence of artifacts. | Blanda | TRUE | DEL
495 | 28 | BA-DEL-001 | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. This step helps detect errors from previous stages, such as treatment prescription or anatomical data collection, and may include checks for adequate slice thickness and absence of artifacts. | Blanda | TRUE | DEL
496 | 30 | BA-DEL-001 | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. | Verification during delineation of treatment volumes and organs at risk by the RTO or RTT. This step helps detect errors from previous stages, such as treatment prescription or anatomical data collection, and may include checks for adequate slice thickness and absence of artifacts. | Blanda | TRUE | DEL
497 | 10 | BA-DEL-002 | Review by the RTO of organs at risk delineated by the RTT. | Review by the RTO of organs at risk delineated by the RTT. This review ensures accuracy and allows detection of errors from earlier stages, such as treatment prescription or anatomical data acquisition. | Normal | TRUE | DEL
498 | 32 | BA-DEL-002 | Review by the RTO of organs at risk delineated by the RTT. | Review by the RTO of organs at risk delineated by the RTT. This review ensures accuracy and allows detection of errors from earlier stages, such as treatment prescription or anatomical data acquisition. | Normal | TRUE | DEL
499 | 33 | BA-DEL-002 | Review by the RTO of organs at risk delineated by the RTT. | Review by the RTO of organs at risk delineated by the RTT. This review ensures accuracy and allows detection of errors from earlier stages, such as treatment prescription or anatomical data acquisition. | Normal | TRUE | DEL
500 | 8 | BA-DEV-001 | Checklist including this parameter for systematic verification. | Verification using checklists or verification sheets. These structured formats are designed to sequentially and systematically monitor compliance with a list of requirements, ensuring that critical parameters are verified thoroughly before completing key actions (e.g., approving a plan or initiating patient irradiation). | Blanda | TRUE | DEV
501 | 10 | BA-DEV-001 | Checklist including this parameter for systematic verification. | Verification using checklists or verification sheets. These structured formats are designed to sequentially and systematically monitor compliance with a list of requirements, ensuring that critical parameters are verified thoroughly before completing key actions (e.g., approving a plan or initiating patient irradiation). | Blanda | TRUE | DEV
502 | 12 | BA-DEV-001 | Checklist including this parameter for systematic verification. | Verification using checklists or verification sheets. These structured formats are designed to sequentially and systematically monitor compliance with a list of requirements, ensuring that critical parameters are verified thoroughly before completing key actions (e.g., approving a plan or initiating patient irradiation). | Blanda | TRUE | DEV
503 | 13 | BA-DEV-001 | Checklist including this parameter for systematic verification. | Verification using checklists or verification sheets. These structured formats are designed to sequentially and systematically monitor compliance with a list of requirements, ensuring that critical parameters are verified thoroughly before completing key actions (e.g., approving a plan or initiating patient irradiation). | Blanda | TRUE | DEV
504 | 14 | BA-DEV-001 | Checklist including this parameter for systematic verification. | Verification using checklists or verification sheets. These structured formats are designed to sequentially and systematically monitor compliance with a list of requirements, ensuring that critical parameters are verified thoroughly before completing key actions (e.g., approving a plan or initiating patient irradiation). | Blanda | TRUE | DEV
505 | 15 | BA-DEV-001 | Checklist including this parameter for systematic verification. | Verification using checklists or verification sheets. These structured formats are designed to sequentially and systematically monitor compliance with a list of requirements, ensuring that critical parameters are verified thoroughly before completing key actions (e.g., approving a plan or initiating patient irradiation). | Blanda | TRUE | DEV
506 | 23 | BA-DEV-001 | Checklist including this parameter for systematic verification. | Verification using checklists or verification sheets. These structured formats are designed to sequentially and systematically monitor compliance with a list of requirements, ensuring that critical parameters are verified thoroughly before completing key actions (e.g., approving a plan or initiating patient irradiation). | Blanda | TRUE | DEV
507 | 29 | BA-DEV-001 | Checklist including this parameter for systematic verification. | Verification using checklists or verification sheets. These structured formats are designed to sequentially and systematically monitor compliance with a list of requirements, ensuring that critical parameters are verified thoroughly before completing key actions (e.g., approving a plan or initiating patient irradiation). | Blanda | TRUE | DEV
508 | 33 | BA-DEV-001 | Checklist including this parameter for systematic verification. | Verification using checklists or verification sheets. These structured formats are designed to sequentially and systematically monitor compliance with a list of requirements, ensuring that critical parameters are verified thoroughly before completing key actions (e.g., approving a plan or initiating patient irradiation). | Blanda | TRUE | DEV
509 | 39 | BA-DEV-001 | Checklist including this parameter for systematic verification. | Verification using checklists or verification sheets. These structured formats are designed to sequentially and systematically monitor compliance with a list of requirements, ensuring that critical parameters are verified thoroughly before completing key actions (e.g., approving a plan or initiating patient irradiation). | Blanda | TRUE | DEV
510 | 41 | BA-DEV-001 | Checklist including this parameter for systematic verification. | Verification using checklists or verification sheets. These structured formats are designed to sequentially and systematically monitor compliance with a list of requirements, ensuring that critical parameters are verified thoroughly before completing key actions (e.g., approving a plan or initiating patient irradiation). | Blanda | TRUE | DEV
511 | 42 | BA-DEV-001 | Checklist including this parameter for systematic verification. | Verification using checklists or verification sheets. These structured formats are designed to sequentially and systematically monitor compliance with a list of requirements, ensuring that critical parameters are verified thoroughly before completing key actions (e.g., approving a plan or initiating patient irradiation). | Blanda | TRUE | DEV
512 | 43 | BA-DEV-001 | Checklist including this parameter for systematic verification. | Verification using checklists or verification sheets. These structured formats are designed to sequentially and systematically monitor compliance with a list of requirements, ensuring that critical parameters are verified thoroughly before completing key actions (e.g., approving a plan or initiating patient irradiation). | Blanda | TRUE | DEV
513 | 44 | BA-DEV-001 | Checklist including this parameter for systematic verification. | Verification using checklists or verification sheets. These structured formats are designed to sequentially and systematically monitor compliance with a list of requirements, ensuring that critical parameters are verified thoroughly before completing key actions (e.g., approving a plan or initiating patient irradiation). | Blanda | TRUE | DEV
514 | 45 | BA-DEV-001 | Checklist including this parameter for systematic verification. | Verification using checklists or verification sheets. These structured formats are designed to sequentially and systematically monitor compliance with a list of requirements, ensuring that critical parameters are verified thoroughly before completing key actions (e.g., approving a plan or initiating patient irradiation). | Blanda | TRUE | DEV
515 | 46 | BA-DEV-001 | Checklist including this parameter for systematic verification. | Verification using checklists or verification sheets. These structured formats are designed to sequentially and systematically monitor compliance with a list of requirements, ensuring that critical parameters are verified thoroughly before completing key actions (e.g., approving a plan or initiating patient irradiation). | Blanda | TRUE | DEV
516 | 47 | BA-DEV-001 | Checklist including this parameter for systematic verification. | Verification using checklists or verification sheets. These structured formats are designed to sequentially and systematically monitor compliance with a list of requirements, ensuring that critical parameters are verified thoroughly before completing key actions (e.g., approving a plan or initiating patient irradiation). | Blanda | TRUE | DEV
517 | 48 | BA-DEV-001 | Checklist including this parameter for systematic verification. | Verification using checklists or verification sheets. These structured formats are designed to sequentially and systematically monitor compliance with a list of requirements, ensuring that critical parameters are verified thoroughly before completing key actions (e.g., approving a plan or initiating patient irradiation). | Blanda | TRUE | DEV
518 | 50 | BA-DEV-001 | Checklist including this parameter for systematic verification. | Verification using checklists or verification sheets. These structured formats are designed to sequentially and systematically monitor compliance with a list of requirements, ensuring that critical parameters are verified thoroughly before completing key actions (e.g., approving a plan or initiating patient irradiation). | Blanda | TRUE | DEV
519 | 73 | BA-DEV-001 | Checklist including this parameter for systematic verification. | Verification using checklists or verification sheets. These structured formats are designed to sequentially and systematically monitor compliance with a list of requirements, ensuring that critical parameters are verified thoroughly before completing key actions (e.g., approving a plan or initiating patient irradiation). | Blanda | TRUE | DEV
520 | 79 | BA-DEV-001 | Checklist including this parameter for systematic verification. | Verification using checklists or verification sheets. These structured formats are designed to sequentially and systematically monitor compliance with a list of requirements, ensuring that critical parameters are verified thoroughly before completing key actions (e.g., approving a plan or initiating patient irradiation). | Blanda | TRUE | DEV
521 | 89 | BA-DEV-001 | Checklist including this parameter for systematic verification. | Verification using checklists or verification sheets. These structured formats are designed to sequentially and systematically monitor compliance with a list of requirements, ensuring that critical parameters are verified thoroughly before completing key actions (e.g., approving a plan or initiating patient irradiation). | Blanda | TRUE | DEV
522 | 90 | BA-DEV-001 | Checklist including this parameter for systematic verification. | Verification using checklists or verification sheets. These structured formats are designed to sequentially and systematically monitor compliance with a list of requirements, ensuring that critical parameters are verified thoroughly before completing key actions (e.g., approving a plan or initiating patient irradiation). | Blanda | TRUE | DEV
523 | 94 | BA-DEV-001 | Checklist including this parameter for systematic verification. | Verification using checklists or verification sheets. These structured formats are designed to sequentially and systematically monitor compliance with a list of requirements, ensuring that critical parameters are verified thoroughly before completing key actions (e.g., approving a plan or initiating patient irradiation). | Blanda | TRUE | DEV
524 | 103 | BA-DEV-001 | Checklist including this parameter for systematic verification. | Verification using checklists or verification sheets. These structured formats are designed to sequentially and systematically monitor compliance with a list of requirements, ensuring that critical parameters are verified thoroughly before completing key actions (e.g., approving a plan or initiating patient irradiation). | Blanda | TRUE | DEV
525 | 58 | BA-DEV-002 | Weekly medical check-up of the patient to monitor treatment progress and overall health. | Weekly medical review of the patient to detect errors in treatment administration or in earlier stages of the process. | Blanda | TRUE | DEV
526 | 104 | BA-DEV-002 | Weekly medical check-up of the patient to monitor treatment progress and overall health. | Weekly medical review of the patient to detect errors in treatment administration or in earlier stages of the process. | Blanda | TRUE | DEV
527 | 105 | BA-DEV-002 | Weekly medical check-up of the patient to monitor treatment progress and overall health. | Weekly medical review of the patient to detect errors in treatment administration or in earlier stages of the process. | Blanda | TRUE | DEV
528 | 8 | BA-DEV-003 | Weekly and end-of-treatment review of the treatment chart ('chart check') by the medical physicist. | Weekly and end-of-treatment review of the treatment chart ('chart check') by the medical physicist to verify that the treatment is being administered as planned. | Blanda | TRUE | DEV
529 | 11 | BA-DEV-003 | Weekly and end-of-treatment review of the treatment chart ('chart check') by the medical physicist. | Weekly and end-of-treatment review of the treatment chart ('chart check') by the medical physicist to verify that the treatment is being administered as planned. | Blanda | TRUE | DEV
530 | 57 | BA-DEV-003 | Weekly and end-of-treatment review of the treatment chart ('chart check') by the medical physicist. | Weekly and end-of-treatment review of the treatment chart ('chart check') by the medical physicist to verify that the treatment is being administered as planned. | Blanda | TRUE | DEV
531 | 104 | BA-DEV-003 | Weekly and end-of-treatment review of the treatment chart ('chart check') by the medical physicist. | Weekly and end-of-treatment review of the treatment chart ('chart check') by the medical physicist to verify that the treatment is being administered as planned. | Blanda | TRUE | DEV
532 | 105 | BA-DEV-003 | Weekly and end-of-treatment review of the treatment chart ('chart check') by the medical physicist. | Weekly and end-of-treatment review of the treatment chart ('chart check') by the medical physicist to verify that the treatment is being administered as planned. | Blanda | TRUE | DEV
533 | 107 | BA-DEV-003 | Weekly and end-of-treatment review of the treatment chart ('chart check') by the medical physicist. | Weekly and end-of-treatment review of the treatment chart ('chart check') by the medical physicist to verify that the treatment is being administered as planned. | Blanda | TRUE | DEV
534 | 53 | BA-DEV-004 | Scheduling a new appointment for the patient to restart treatment. | Schedule a new appointment for the patient to restart treatment (e.g., issuing a new initial treatment session card) and cancel any continuation appointments to prevent treatment if plan changes require replanning. | Blanda | TRUE | DEV
535 | 2 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
536 | 3 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
537 | 4 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
538 | 5 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
539 | 6 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
540 | 7 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
541 | 8 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
542 | 10 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
543 | 11 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
544 | 12 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
545 | 17 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
546 | 18 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
547 | 22 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
548 | 23 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
549 | 28 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
550 | 30 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
551 | 31 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
552 | 32 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
553 | 33 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
554 | 34 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
555 | 36 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
556 | 37 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
557 | 39 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
558 | 40 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
559 | 47 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
560 | 50 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
561 | 51 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
562 | 52 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
563 | 54 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
564 | 55 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
565 | 59 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
566 | 82 | BA-DEV-005 | Discussion of cases prior to the first session of treatment initiation during a technical/clinical session, including any plan modifications. | Presentation and discussion of cases prior to the start of treatment in a daily technical/clinical session, with RTO, QMP, and RTT from the facility itself (Internal Peer Review), in which various aspects of the treatment are discussed, including possible changes to the plan. | Normal | TRUE | DEV
567 | 60 | BA-DEV-006 | Alternative system for recording the stage of the radiotherapy process for each patient. | Alternative registration system (e.g., card, Excel sheet) to record the stage of the radiotherapy process for each patient. | Blanda | TRUE | DEV
568 | 9 | BA-IDE-001 | Manual verification of the patient’s identity using a photograph, ID card, medical record number, or similar identifiers. | Manual verification of patient identity using a photograph taken at an earlier stage (preferably during medical record creation), ID number, or medical record number to detect identification errors during the treatment process. | Blanda | TRUE | IDE
569 | 20 | BA-IDE-001 | Manual verification of the patient’s identity using a photograph, ID card, medical record number, or similar identifiers. | Manual verification of patient identity using a photograph taken at an earlier stage (preferably during medical record creation), ID number, or medical record number to detect identification errors during the treatment process. | Blanda | TRUE | IDE
570 | 34 | BA-IDE-001 | Manual verification of the patient’s identity using a photograph, ID card, medical record number, or similar identifiers. | Manual verification of patient identity using a photograph taken at an earlier stage (preferably during medical record creation), ID number, or medical record number to detect identification errors during the treatment process. | Blanda | TRUE | IDE
571 | 65 | BA-IDE-001 | Manual verification of the patient’s identity using a photograph, ID card, medical record number, or similar identifiers. | Manual verification of patient identity using a photograph taken at an earlier stage (preferably during medical record creation), ID number, or medical record number to detect identification errors during the treatment process. | Blanda | TRUE | IDE
572 | 1 | BA-IDE-002 | Verification of the patient's identity by asking the patient their first name, last name, date of birth, and the area to be treated. | Verification of the patient's identity by asking the patient their first name, last name, date of birth, and the area to be treated. | Blanda | TRUE | IDE
573 | 9 | BA-IDE-002 | Verification of the patient's identity by asking the patient their first name, last name, date of birth, and the area to be treated. | Verification of the patient's identity by asking the patient their first name, last name, date of birth, and the area to be treated. | Blanda | TRUE | IDE
574 | 20 | BA-IDE-002 | Verification of the patient's identity by asking the patient their first name, last name, date of birth, and the area to be treated. | Verification of the patient's identity by asking the patient their first name, last name, date of birth, and the area to be treated. | Blanda | TRUE | IDE
575 | 64 | BA-IDE-002 | Verification of the patient's identity by asking the patient their first name, last name, date of birth, and the area to be treated. | Verification of the patient's identity by asking the patient their first name, last name, date of birth, and the area to be treated. | Blanda | TRUE | IDE
576 | 65 | BA-IDE-002 | Verification of the patient's identity by asking the patient their first name, last name, date of birth, and the area to be treated. | Verification of the patient's identity by asking the patient their first name, last name, date of birth, and the area to be treated. | Blanda | TRUE | IDE
577 | 83 | BA-IDE-002 | Verification of the patient's identity by asking the patient their first name, last name, date of birth, and the area to be treated. | Verification of the patient's identity by asking the patient their first name, last name, date of birth, and the area to be treated. | Blanda | TRUE | IDE
578 | 84 | BA-IDE-002 | Verify the patient's identity by asking for their first name, last name, date of birth, and the area to be treated. | Verify the patient's identity by asking for their first name, last name, date of birth, and the area to be treated. | Blanda | TRUE | IDE
579 | 9 | BA-IDE-003 | Verify the patient's identity using electronic identification systems. | Verify the patient's identity using electronic identification systems (e.g., chip reading, barcode scanning, fingerprint scanning, facial recognition) to prevent identification errors during the treatment process. | Robusta | TRUE | IDE
580 | 20 | BA-IDE-003 | Verify the patient's identity using electronic identification systems. | Verify the patient's identity using electronic identification systems (e.g., chip reading, barcode scanning, fingerprint scanning, facial recognition) to prevent identification errors during the treatment process. | Robusta | TRUE | IDE
581 | 64 | BA-IDE-003 | Verify the patient's identity using electronic identification systems. | Verify the patient's identity using electronic identification systems (e.g., chip reading, barcode scanning, fingerprint scanning, facial recognition) to prevent identification errors during the treatment process. | Robusta | TRUE | IDE
582 | 65 | BA-IDE-003 | Verify the patient's identity using electronic identification systems. | Verify the patient's identity using electronic identification systems (e.g., chip reading, barcode scanning, fingerprint scanning, facial recognition) to prevent identification errors during the treatment process. | Robusta | TRUE | IDE
583 | 83 | BA-IDE-003 | Verify the patient's identity using electronic identification systems. | Verify the patient's identity using electronic identification systems (e.g., chip reading, barcode scanning, fingerprint scanning, facial recognition) to prevent identification errors during the treatment process. | Robusta | TRUE | IDE
584 | 84 | BA-IDE-003 | Verify the patient's identity using electronic identification systems. | Verify the patient's identity using electronic identification systems (e.g., chip reading, barcode scanning, fingerprint scanning, facial recognition) to prevent identification errors during the treatment process. | Robusta | TRUE | IDE
585 | 9 | BA-LOC-001 | Verification of the image acquisition process by a second RTT. | Independent verification by a second RTT of the patient placement and image acquisition during the simulation process. | Normal | TRUE | LOC
586 | 13 | BA-LOC-001 | Verification of the image acquisition process by a second RTT. | Independent verification by a second RTT of the patient placement and image acquisition during the simulation process. | Normal | TRUE | LOC
587 | 14 | BA-LOC-001 | Verification of the image acquisition process by a second RTT. | Independent verification by a second RTT of the patient placement and image acquisition during the simulation process. | Normal | TRUE | LOC
588 | 15 | BA-LOC-001 | Verification of the image acquisition process by a second RTT. | Independent verification by a second RTT of the patient placement and image acquisition during the simulation process. | Normal | TRUE | LOC
589 | 16 | BA-LOC-001 | Verification of the image acquisition process by a second RTT. | Independent verification by a second RTT of the patient placement and image acquisition during the simulation process. | Normal | TRUE | LOC
590 | 19 | BA-LOC-001 | Verification of the image acquisition process by a second RTT. | Independent verification by a second RTT of the patient placement and image acquisition during the simulation process. | Normal | TRUE | LOC
591 | 21 | BA-LOC-001 | Verification of the image acquisition process by a second RTT. | Independent verification by a second RTT of the patient placement and image acquisition during the simulation process. | Normal | TRUE | LOC
592 | 24 | BA-LOC-001 | Verification of the image acquisition process by a second RTT. | Independent verification by a second RTT of the patient placement and image acquisition during the simulation process. | Normal | TRUE | LOC
593 | 25 | BA-LOC-001 | Verification of the image acquisition process by a second RTT. | Independent verification by a second RTT of the patient placement and image acquisition during the simulation process. | Normal | TRUE | LOC
594 | 26 | BA-LOC-001 | Verification of the image acquisition process by a second RTT. | Independent verification by a second RTT of the patient placement and image acquisition during the simulation process. | Normal | TRUE | LOC
595 | 27 | BA-LOC-001 | Verification of the image acquisition process by a second RTT. | Independent verification by a second RTT of the patient placement and image acquisition during the simulation process. | Normal | TRUE | LOC
596 | 11 | BA-LOC-002 | Verification of the patient's preparation conditions using specific parameters visible in the acquired images. | Verification of the patient's preparation conditions using specific parameters visible in the acquired images. Includes checking organ volumes (e.g., bladder, rectum, stomach) and respiratory phase by measuring distances and volumes and comparing them against established tolerances. | Blanda | TRUE | LOC
597 | 41 | BA-LOC-002 | Verification of the patient's preparation conditions using specific parameters visible in the acquired images. | Verification of the patient's preparation conditions using specific parameters visible in the acquired images. Includes checking organ volumes (e.g., bladder, rectum, stomach) and respiratory phase by measuring distances and volumes and comparing them against established tolerances. | Blanda | TRUE | LOC
598 | 93 | BA-LOC-002 | Verification of the patient's preparation conditions using specific parameters visible in the acquired images. | Verification of the patient's preparation conditions using specific parameters visible in the acquired images. Includes checking organ volumes (e.g., bladder, rectum, stomach) and respiratory phase by measuring distances and volumes and comparing them against established tolerances. | Blanda | TRUE | LOC
599 | 29 | BA-LOC-003 | System alert triggered when patient identification data does not match. | System alert when patient identification data does not match the demographic data associated with the images. | Muy robusta | TRUE | LOC
600 | 30 | BA-LOC-003 | System alert triggered when patient identification data does not match. | System alert when patient identification data does not match the demographic data associated with the images. | Muy robusta | TRUE | LOC
601 | 9 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
602 | 13 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
603 | 14 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
604 | 24 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
605 | 25 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
606 | 26 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
607 | 27 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
608 | 38 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
609 | 41 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
610 | 42 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
611 | 61 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
612 | 64 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
613 | 65 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
614 | 66 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
615 | 67 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
616 | 68 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
617 | 69 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
618 | 70 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
619 | 71 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
620 | 72 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
621 | 73 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
622 | 74 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
623 | 75 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
624 | 76 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
625 | 77 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
626 | 78 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
627 | 79 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
628 | 80 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
629 | 83 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
630 | 84 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
631 | 85 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
632 | 86 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
633 | 87 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
634 | 88 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
635 | 89 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
636 | 90 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
637 | 91 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
638 | 92 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
639 | 94 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
640 | 95 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
641 | 96 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
642 | 97 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
643 | 98 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
644 | 99 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
645 | 101 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
646 | 102 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
647 | 103 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
648 | 110 | BA-LOC-004 | Automatic surface recognition system (SGRT) for surface verification and respiratory cycle monitoring. | Use of external automated systems to verify patient positioning through surface recognition and respiratory cycle monitoring. | Robusta | TRUE | LOC
649 | 10 | BA-LOC-005 | Redundant review of imaging indications and immobilization proposal by another specialist (RTO, QMP, or RTT). | Redundant review of imaging indications and immobilization proposal by another specialist (RTO, QMP, or RTT). | Normal | TRUE | LOC
650 | 9 | BA-LOC-006 | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Normal | TRUE | LOC
651 | 10 | BA-LOC-006 | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Normal | TRUE | LOC
652 | 11 | BA-LOC-006 | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Normal | TRUE | LOC
653 | 12 | BA-LOC-006 | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Normal | TRUE | LOC
654 | 13 | BA-LOC-006 | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Normal | TRUE | LOC
655 | 14 | BA-LOC-006 | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Normal | TRUE | LOC
656 | 15 | BA-LOC-006 | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Normal | TRUE | LOC
657 | 16 | BA-LOC-006 | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Normal | TRUE | LOC
658 | 17 | BA-LOC-006 | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Normal | TRUE | LOC
659 | 18 | BA-LOC-006 | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Normal | TRUE | LOC
660 | 19 | BA-LOC-006 | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Normal | TRUE | LOC
661 | 21 | BA-LOC-006 | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Normal | TRUE | LOC
662 | 22 | BA-LOC-006 | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Normal | TRUE | LOC
663 | 23 | BA-LOC-006 | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Presence of the RTO and, when required, QMP or nursing staff during image acquisition. | Normal | TRUE | LOC
664 | 67 | BA-LOC-007 | Electronic identification systems for immobilization devices. | Use of electronic identification systems for immobilization devices to prevent irradiation if the immobilizer does not correspond to the correct patient. | Robusta | TRUE | LOC
665 | 86 | BA-LOC-007 | Electronic identification systems for immobilization devices. | Use of electronic identification systems for immobilization devices to prevent irradiation if the immobilizer does not correspond to the correct patient. | Robusta | TRUE | LOC
666 | 11 | BA-LOC-008 | Verification of images following CT scan acquisition. | Verification of images following CT scan acquisition. | Blanda | TRUE | LOC
667 | 12 | BA-LOC-008 | Verification of images following CT scan acquisition. | Verification of images following CT scan acquisition. | Blanda | TRUE | LOC
668 | 14 | BA-LOC-008 | Verification of images following CT scan acquisition. | Verification of images following CT scan acquisition. | Blanda | TRUE | LOC
669 | 18 | BA-LOC-008 | Verification of images following CT scan acquisition. | Verification of images following CT scan acquisition. | Blanda | TRUE | LOC
670 | 23 | BA-LOC-009 | Review of images for fusion by a radiologist or nuclear medicine physician. | Prior review of images for fusion by a radiologist or nuclear medicine physician. | Blanda | TRUE | LOC
671 | 49 | BA-MAN-001 | Verification of TPS whenever any modification is made to its operating procedures. | Verification of the TPS required by a service procedure , including checks such as planning validation, comparison with reference data, and measurement of cases after any TPS modifications, to ensure validation and approval for clinical use. | Normal | TRUE | MAN
672 | 2 | BA-PLA-001 | Review of prescription and simulation data prior to treatment planning. | Review by the QMP or dosimetrist prior to planning to ensure prescription data and patient images are correct. | Normal | TRUE | PLA
673 | 3 | BA-PLA-001 | Review of prescription and simulation data prior to treatment planning. | Review by the QMP or dosimetrist prior to planning to ensure prescription data and patient images are correct. | Normal | TRUE | PLA
674 | 4 | BA-PLA-001 | Review of prescription and simulation data prior to treatment planning. | Review by the QMP or dosimetrist prior to planning to ensure prescription data and patient images are correct. | Normal | TRUE | PLA
675 | 5 | BA-PLA-001 | Review of prescription and simulation data prior to treatment planning. | Review by the QMP or dosimetrist prior to planning to ensure prescription data and patient images are correct. | Normal | TRUE | PLA
676 | 10 | BA-PLA-001 | Review of prescription and simulation data prior to treatment planning. | Review by the QMP or dosimetrist prior to planning to ensure prescription data and patient images are correct. | Normal | TRUE | PLA
677 | 11 | BA-PLA-001 | Review of prescription and simulation data prior to treatment planning. | Review by the QMP or dosimetrist prior to planning to ensure prescription data and patient images are correct. | Normal | TRUE | PLA
678 | 12 | BA-PLA-001 | Review of prescription and simulation data prior to treatment planning. | Review by the QMP or dosimetrist prior to planning to ensure prescription data and patient images are correct. | Normal | TRUE | PLA
679 | 13 | BA-PLA-001 | Review of prescription and simulation data prior to treatment planning. | Review by the QMP or dosimetrist prior to planning to ensure prescription data and patient images are correct. | Normal | TRUE | PLA
680 | 14 | BA-PLA-001 | Review of prescription and simulation data prior to treatment planning. | Review by the QMP or dosimetrist prior to planning to ensure prescription data and patient images are correct. | Normal | TRUE | PLA
681 | 15 | BA-PLA-001 | Review of prescription and simulation data prior to treatment planning. | Review by the QMP or dosimetrist prior to planning to ensure prescription data and patient images are correct. | Normal | TRUE | PLA
682 | 17 | BA-PLA-001 | Review of prescription and simulation data prior to treatment planning. | Review by the QMP or dosimetrist prior to planning to ensure prescription data and patient images are correct. | Normal | TRUE | PLA
683 | 18 | BA-PLA-001 | Review of prescription and simulation data prior to treatment planning. | Review by the QMP or dosimetrist prior to planning to ensure prescription data and patient images are correct. | Normal | TRUE | PLA
684 | 19 | BA-PLA-001 | Review of prescription and simulation data prior to treatment planning. | Review by the QMP or dosimetrist prior to planning to ensure prescription data and patient images are correct. | Normal | TRUE | PLA
685 | 20 | BA-PLA-001 | Review of prescription and simulation data prior to treatment planning. | Review by the QMP or dosimetrist prior to planning to ensure prescription data and patient images are correct. | Normal | TRUE | PLA
686 | 21 | BA-PLA-001 | Review of prescription and simulation data prior to treatment planning. | Review by the QMP or dosimetrist prior to planning to ensure prescription data and patient images are correct. | Normal | TRUE | PLA
687 | 26 | BA-PLA-001 | Review of prescription and simulation data prior to treatment planning. | Review by the QMP or dosimetrist prior to planning to ensure prescription data and patient images are correct. | Normal | TRUE | PLA
688 | 28 | BA-PLA-001 | Review of prescription and simulation data prior to treatment planning. | Review by the QMP or dosimetrist prior to planning to ensure prescription data and patient images are correct. | Normal | TRUE | PLA
689 | 30 | BA-PLA-001 | Review of prescription and simulation data prior to treatment planning. | Review by the QMP or dosimetrist prior to planning to ensure prescription data and patient images are correct. | Normal | TRUE | PLA
690 | 31 | BA-PLA-001 | Review of prescription and simulation data prior to treatment planning. | Review by the QMP or dosimetrist prior to planning to ensure prescription data and patient images are correct. | Normal | TRUE | PLA
691 | 32 | BA-PLA-001 | Review of prescription and simulation data prior to treatment planning. | Review by the QMP or dosimetrist prior to planning to ensure prescription data and patient images are correct. | Normal | TRUE | PLA
692 | 33 | BA-PLA-001 | Review of prescription and simulation data prior to treatment planning. | Review by the QMP or dosimetrist prior to planning to ensure prescription data and patient images are correct. | Normal | TRUE | PLA
693 | 35 | BA-PLA-001 | Review of prescription and simulation data prior to treatment planning. | Review by the QMP or dosimetrist prior to planning to ensure prescription data and patient images are correct. | Normal | TRUE | PLA
694 | 36 | BA-PLA-001 | Review of prescription and simulation data prior to treatment planning. | Review by the QMP or dosimetrist prior to planning to ensure prescription data and patient images are correct. | Normal | TRUE | PLA
695 | 37 | BA-PLA-001 | Review of prescription and simulation data prior to treatment planning. | Review by the QMP or dosimetrist prior to planning to ensure prescription data and patient images are correct. | Normal | TRUE | PLA
696 | 40 | BA-PLA-001 | Review of prescription and simulation data prior to treatment planning. | Review by the QMP or dosimetrist prior to planning to ensure prescription data and patient images are correct. | Normal | TRUE | PLA
697 | 44 | BA-PLA-002 | Independent verification of dosimetric calculations. | Verification of planning results through independent dosimetric calculation using a separate tool to confirm the accuracy of TPS data and models. | Normal | TRUE | PLA
698 | 45 | BA-PLA-002 | Independent verification of dosimetric calculations. | Verification of planning results through independent dosimetric calculation using a separate tool to confirm the accuracy of TPS data and models. | Normal | TRUE | PLA
699 | 47 | BA-PLA-002 | Independent verification of dosimetric calculations. | Verification of planning results through independent dosimetric calculation using a separate tool to confirm the accuracy of TPS data and models. | Normal | TRUE | PLA
700 | 49 | BA-PLA-002 | Independent verification of dosimetric calculations. | Verification of planning results through independent dosimetric calculation using a separate tool to confirm the accuracy of TPS data and models. | Normal | TRUE | PLA
701 | 52 | BA-PLA-002 | Independent verification of dosimetric calculations. | Verification of planning results through independent dosimetric calculation using a separate tool to confirm the accuracy of TPS data and models. | Normal | TRUE | PLA
702 | 59 | BA-PLA-002 | Independent verification of dosimetric calculations. | Verification of planning results through independent dosimetric calculation using a separate tool to confirm the accuracy of TPS data and models. | Normal | TRUE | PLA
703 | 82 | BA-PLA-002 | Independent verification of dosimetric calculations. | Verification of planning results through independent dosimetric calculation using a separate tool to confirm the accuracy of TPS data and models. | Normal | TRUE | PLA
704 | 2 | BA-PLA-003 | Independent review of the treatment plan by a QMP. | Verification of treatment dosimetry planning results by a QMP other than the planner, using different methodologies or automated scripts to check various parameters. | Normal | TRUE | PLA
705 | 3 | BA-PLA-003 | Independent review of the treatment plan by a QMP. | Verification of treatment dosimetry planning results by a QMP other than the planner, using different methodologies or automated scripts to check various parameters. | Normal | TRUE | PLA
706 | 4 | BA-PLA-003 | Independent review of the treatment plan by a QMP. | Verification of treatment dosimetry planning results by a QMP other than the planner, using different methodologies or automated scripts to check various parameters. | Normal | TRUE | PLA
707 | 39 | BA-PLA-003 | Independent review of the treatment plan by a QMP. | Verification of treatment dosimetry planning results by a QMP other than the planner, using different methodologies or automated scripts to check various parameters. | Normal | TRUE | PLA
708 | 41 | BA-PLA-003 | Independent review of the treatment plan by a QMP. | Verification of treatment dosimetry planning results by a QMP other than the planner, using different methodologies or automated scripts to check various parameters. | Normal | TRUE | PLA
709 | 42 | BA-PLA-003 | Independent review of the treatment plan by a QMP. | Verification of treatment dosimetry planning results by a QMP other than the planner, using different methodologies or automated scripts to check various parameters. | Normal | TRUE | PLA
710 | 43 | BA-PLA-003 | Independent review of the treatment plan by a QMP. | Verification of treatment dosimetry planning results by a QMP other than the planner, using different methodologies or automated scripts to check various parameters. | Normal | TRUE | PLA
711 | 44 | BA-PLA-003 | Independent review of the treatment plan by a QMP. | Verification of treatment dosimetry planning results by a QMP other than the planner, using different methodologies or automated scripts to check various parameters. | Normal | TRUE | PLA
712 | 45 | BA-PLA-003 | Independent review of the treatment plan by a QMP. | Verification of treatment dosimetry planning results by a QMP other than the planner, using different methodologies or automated scripts to check various parameters. | Normal | TRUE | PLA
713 | 46 | BA-PLA-003 | Independent review of the treatment plan by a QMP. | Verification of treatment dosimetry planning results by a QMP other than the planner, using different methodologies or automated scripts to check various parameters. | Normal | TRUE | PLA
714 | 47 | BA-PLA-003 | Independent review of the treatment plan by a QMP. | Verification of treatment dosimetry planning results by a QMP other than the planner, using different methodologies or automated scripts to check various parameters. | Normal | TRUE | PLA
715 | 48 | BA-PLA-003 | Independent review of the treatment plan by a QMP. | Verification of treatment dosimetry planning results by a QMP other than the planner, using different methodologies or automated scripts to check various parameters. | Normal | TRUE | PLA
716 | 49 | BA-PLA-003 | Independent review of the treatment plan by a QMP. | Verification of treatment dosimetry planning results by a QMP other than the planner, using different methodologies or automated scripts to check various parameters. | Normal | TRUE | PLA
717 | 50 | BA-PLA-003 | Independent review of the treatment plan by a QMP. | Verification of treatment dosimetry planning results by a QMP other than the planner, using different methodologies or automated scripts to check various parameters. | Normal | TRUE | PLA
718 | 51 | BA-PLA-003 | Independent review of the treatment plan by a QMP. | Verification of treatment dosimetry planning results by a QMP other than the planner, using different methodologies or automated scripts to check various parameters. | Normal | TRUE | PLA
719 | 54 | BA-PLA-003 | Independent review of the treatment plan by a QMP. | Verification of treatment dosimetry planning results by a QMP other than the planner, using different methodologies or automated scripts to check various parameters. | Normal | TRUE | PLA
720 | 1 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
721 | 2 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
722 | 3 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
723 | 4 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
724 | 5 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
725 | 6 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
726 | 7 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
727 | 8 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
728 | 11 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
729 | 12 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
730 | 15 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
731 | 17 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
732 | 18 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
733 | 19 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
734 | 29 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
735 | 32 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
736 | 33 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
737 | 34 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
738 | 35 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
739 | 36 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
740 | 37 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
741 | 38 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
742 | 39 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
743 | 41 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
744 | 42 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
745 | 47 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
746 | 49 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
747 | 50 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
748 | 51 | BA-PLA-004 | Joint evaluation of the dosimetry plan by the RTO and the QMP. | Joint evaluation of the dosimetry plan by the RTO and QMP during the planning stage that allows to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
749 | 40 | BA-PLA-005 | Joint evaluation of the new dosimetry plan by the RTO and the QMP. | Joint evaluation of the NEW dosimetry plan by the RTO and QMP during the planning stage to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
750 | 53 | BA-PLA-005 | Joint evaluation of the new dosimetry plan by the RTO and the QMP. | Joint evaluation of the NEW dosimetry plan by the RTO and QMP during the planning stage to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
751 | 54 | BA-PLA-005 | Joint evaluation of the new dosimetry plan by the RTO and the QMP. | Joint evaluation of the NEW dosimetry plan by the RTO and QMP during the planning stage to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
752 | 82 | BA-PLA-005 | Joint evaluation of the new dosimetry plan by the RTO and the QMP. | Joint evaluation of the NEW dosimetry plan by the RTO and QMP during the planning stage to detect potential dosimetric or geometric errors from previous steps. | Normal | TRUE | PLA
753 | 46 | BA-PLA-006 | Interlock system that prevents treatment on a machine other than the one assigned in the treatment plan. | Interlock system that prevents treatment on a machine other than the one assigned in the treatment plan. | Muy robusta | TRUE | PLA
754 | 38 | BA-PLA-007 | Software notification indicating a discrepancy between demographic data in the DICOM header and patient data in the treatment plan. | Software notification indicating a discrepancy between demographic data in the DICOM header and patient data in the treatment plan. | Robusta | TRUE | PLA
755 | 9 | BA-SIN-001 | Positioning and immobilization of the patient in the treatment position for special sessions (e.g., initial session), in the presence of the RTO and RTT, and whenever possible, the QMP. | Positioning and immobilization of the patient in the presence of the RTO and RTT, and whenever possible the QMP, following a procedure that defines the tasks of each professional involved. This step enables detection of geometric, dosimetric, or patient identification errors from previous stages of the treatment planning process. The barrier applies whenever the error can be identified through the functions performed by these professionals and is not explicitly covered by another barrier assigned to the same initiating event. | Normal | TRUE | SIN
756 | 13 | BA-SIN-001 | Positioning and immobilization of the patient in the treatment position for special sessions (e.g., initial session), in the presence of the RTO and RTT, and whenever possible, the QMP. | Positioning and immobilization of the patient in the presence of the RTO and RTT, and whenever possible the QMP, following a procedure that defines the tasks of each professional involved. This step enables detection of geometric, dosimetric, or patient identification errors from previous stages of the treatment planning process. The barrier applies whenever the error can be identified through the functions performed by these professionals and is not explicitly covered by another barrier assigned to the same initiating event. | Normal | TRUE | SIN
757 | 15 | BA-SIN-001 | Positioning and immobilization of the patient in the treatment position for special sessions (e.g., initial session), in the presence of the RTO and RTT, and whenever possible, the QMP. | Positioning and immobilization of the patient in the presence of the RTO and RTT, and whenever possible the QMP, following a procedure that defines the tasks of each professional involved. This step enables detection of geometric, dosimetric, or patient identification errors from previous stages of the treatment planning process. The barrier applies whenever the error can be identified through the functions performed by these professionals and is not explicitly covered by another barrier assigned to the same initiating event. | Normal | TRUE | SIN
758 | 16 | BA-SIN-001 | Positioning and immobilization of the patient in the treatment position for special sessions (e.g., initial session), in the presence of the RTO and RTT, and whenever possible, the QMP. | Positioning and immobilization of the patient in the presence of the RTO and RTT, and whenever possible the QMP, following a procedure that defines the tasks of each professional involved. This step enables detection of geometric, dosimetric, or patient identification errors from previous stages of the treatment planning process. The barrier applies whenever the error can be identified through the functions performed by these professionals and is not explicitly covered by another barrier assigned to the same initiating event. | Normal | TRUE | SIN
759 | 25 | BA-SIN-001 | Positioning and immobilization of the patient in the treatment position for special sessions (e.g., initial session), in the presence of the RTO and RTT, and whenever possible, the QMP. | Positioning and immobilization of the patient in the presence of the RTO and RTT, and whenever possible the QMP, following a procedure that defines the tasks of each professional involved. This step enables detection of geometric, dosimetric, or patient identification errors from previous stages of the treatment planning process. The barrier applies whenever the error can be identified through the functions performed by these professionals and is not explicitly covered by another barrier assigned to the same initiating event. | Normal | TRUE | SIN
760 | 34 | BA-SIN-001 | Positioning and immobilization of the patient in the treatment position for special sessions (e.g., initial session), in the presence of the RTO and RTT, and whenever possible, the QMP. | Positioning and immobilization of the patient in the presence of the RTO and RTT, and whenever possible the QMP, following a procedure that defines the tasks of each professional involved. This step enables detection of geometric, dosimetric, or patient identification errors from previous stages of the treatment planning process. The barrier applies whenever the error can be identified through the functions performed by these professionals and is not explicitly covered by another barrier assigned to the same initiating event. | Normal | TRUE | SIN
761 | 36 | BA-SIN-001 | Positioning and immobilization of the patient in the treatment position for special sessions (e.g., initial session), in the presence of the RTO and RTT, and whenever possible, the QMP. | Positioning and immobilization of the patient in the presence of the RTO and RTT, and whenever possible the QMP, following a procedure that defines the tasks of each professional involved. This step enables detection of geometric, dosimetric, or patient identification errors from previous stages of the treatment planning process. The barrier applies whenever the error can be identified through the functions performed by these professionals and is not explicitly covered by another barrier assigned to the same initiating event. | Normal | TRUE | SIN
762 | 42 | BA-SIN-001 | Positioning and immobilization of the patient in the treatment position for special sessions (e.g., initial session), in the presence of the RTO and RTT, and whenever possible, the QMP. | Positioning and immobilization of the patient in the presence of the RTO and RTT, and whenever possible the QMP, following a procedure that defines the tasks of each professional involved. This step enables detection of geometric, dosimetric, or patient identification errors from previous stages of the treatment planning process. The barrier applies whenever the error can be identified through the functions performed by these professionals and is not explicitly covered by another barrier assigned to the same initiating event. | Normal | TRUE | SIN
763 | 50 | BA-SIN-001 | Positioning and immobilization of the patient in the treatment position for special sessions (e.g., initial session), in the presence of the RTO and RTT, and whenever possible, the QMP. | Positioning and immobilization of the patient in the presence of the RTO and RTT, and whenever possible the QMP, following a procedure that defines the tasks of each professional involved. This step enables detection of geometric, dosimetric, or patient identification errors from previous stages of the treatment planning process. The barrier applies whenever the error can be identified through the functions performed by these professionals and is not explicitly covered by another barrier assigned to the same initiating event. | Normal | TRUE | SIN
764 | 56 | BA-SIN-001 | Positioning and immobilization of the patient in the treatment position for special sessions (e.g., initial session), in the presence of the RTO and RTT, and whenever possible, the QMP. | Positioning and immobilization of the patient in the presence of the RTO and RTT, and whenever possible the QMP, following a procedure that defines the tasks of each professional involved. This step enables detection of geometric, dosimetric, or patient identification errors from previous stages of the treatment planning process. The barrier applies whenever the error can be identified through the functions performed by these professionals and is not explicitly covered by another barrier assigned to the same initiating event. | Normal | TRUE | SIN
765 | 57 | BA-SIN-001 | Positioning and immobilization of the patient in the treatment position for special sessions (e.g., initial session), in the presence of the RTO and RTT, and whenever possible, the QMP. | Positioning and immobilization of the patient in the presence of the RTO and RTT, and whenever possible the QMP, following a procedure that defines the tasks of each professional involved. This step enables detection of geometric, dosimetric, or patient identification errors from previous stages of the treatment planning process. The barrier applies whenever the error can be identified through the functions performed by these professionals and is not explicitly covered by another barrier assigned to the same initiating event. | Normal | TRUE | SIN
766 | 61 | BA-SIN-001 | Positioning and immobilization of the patient in the treatment position for special sessions (e.g., initial session), in the presence of the RTO and RTT, and whenever possible, the QMP. | Positioning and immobilization of the patient in the presence of the RTO and RTT, and whenever possible the QMP, following a procedure that defines the tasks of each professional involved. This step enables detection of geometric, dosimetric, or patient identification errors from previous stages of the treatment planning process. The barrier applies whenever the error can be identified through the functions performed by these professionals and is not explicitly covered by another barrier assigned to the same initiating event. | Normal | TRUE | SIN
767 | 64 | BA-SIN-001 | Positioning and immobilization of the patient in the treatment position for special sessions (e.g., initial session), in the presence of the RTO and RTT, and whenever possible, the QMP. | Positioning and immobilization of the patient in the presence of the RTO and RTT, and whenever possible the QMP, following a procedure that defines the tasks of each professional involved. This step enables detection of geometric, dosimetric, or patient identification errors from previous stages of the treatment planning process. The barrier applies whenever the error can be identified through the functions performed by these professionals and is not explicitly covered by another barrier assigned to the same initiating event. | Normal | TRUE | SIN
768 | 65 | BA-SIN-001 | Positioning and immobilization of the patient in the treatment position for special sessions (e.g., initial session), in the presence of the RTO and RTT, and whenever possible, the QMP. | Positioning and immobilization of the patient in the presence of the RTO and RTT, and whenever possible the QMP, following a procedure that defines the tasks of each professional involved. This step enables detection of geometric, dosimetric, or patient identification errors from previous stages of the treatment planning process. The barrier applies whenever the error can be identified through the functions performed by these professionals and is not explicitly covered by another barrier assigned to the same initiating event. | Normal | TRUE | SIN
769 | 66 | BA-SIN-001 | Positioning and immobilization of the patient in the treatment position for special sessions (e.g., initial session), in the presence of the RTO and RTT, and whenever possible, the QMP. | Positioning and immobilization of the patient in the presence of the RTO and RTT, and whenever possible the QMP, following a procedure that defines the tasks of each professional involved. This step enables detection of geometric, dosimetric, or patient identification errors from previous stages of the treatment planning process. The barrier applies whenever the error can be identified through the functions performed by these professionals and is not explicitly covered by another barrier assigned to the same initiating event. | Normal | TRUE | SIN
770 | 67 | BA-SIN-001 | Positioning and immobilization of the patient in the treatment position for special sessions (e.g., initial session), in the presence of the RTO and RTT, and whenever possible, the QMP. | Positioning and immobilization of the patient in the presence of the RTO and RTT, and whenever possible the QMP, following a procedure that defines the tasks of each professional involved. This step enables detection of geometric, dosimetric, or patient identification errors from previous stages of the treatment planning process. The barrier applies whenever the error can be identified through the functions performed by these professionals and is not explicitly covered by another barrier assigned to the same initiating event. | Normal | TRUE | SIN
771 | 69 | BA-SIN-001 | Positioning and immobilization of the patient in the treatment position for special sessions (e.g., initial session), in the presence of the RTO and RTT, and whenever possible, the QMP. | Positioning and immobilization of the patient in the presence of the RTO and RTT, and whenever possible the QMP, following a procedure that defines the tasks of each professional involved. This step enables detection of geometric, dosimetric, or patient identification errors from previous stages of the treatment planning process. The barrier applies whenever the error can be identified through the functions performed by these professionals and is not explicitly covered by another barrier assigned to the same initiating event. | Normal | TRUE | SIN
772 | 72 | BA-SIN-001 | Positioning and immobilization of the patient in the treatment position for special sessions (e.g., initial session), in the presence of the RTO and RTT, and whenever possible, the QMP. | Positioning and immobilization of the patient in the presence of the RTO and RTT, and whenever possible the QMP, following a procedure that defines the tasks of each professional involved. This step enables detection of geometric, dosimetric, or patient identification errors from previous stages of the treatment planning process. The barrier applies whenever the error can be identified through the functions performed by these professionals and is not explicitly covered by another barrier assigned to the same initiating event. | Normal | TRUE | SIN
773 | 73 | BA-SIN-001 | Positioning and immobilization of the patient in the treatment position for special sessions (e.g., initial session), in the presence of the RTO and RTT, and whenever possible, the QMP. | Positioning and immobilization of the patient in the presence of the RTO and RTT, and whenever possible the QMP, following a procedure that defines the tasks of each professional involved. This step enables detection of geometric, dosimetric, or patient identification errors from previous stages of the treatment planning process. The barrier applies whenever the error can be identified through the functions performed by these professionals and is not explicitly covered by another barrier assigned to the same initiating event. | Normal | TRUE | SIN
774 | 76 | BA-SIN-001 | Positioning and immobilization of the patient in the treatment position for special sessions (e.g., initial session), in the presence of the RTO and RTT, and whenever possible, the QMP. | Positioning and immobilization of the patient in the presence of the RTO and RTT, and whenever possible the QMP, following a procedure that defines the tasks of each professional involved. This step enables detection of geometric, dosimetric, or patient identification errors from previous stages of the treatment planning process. The barrier applies whenever the error can be identified through the functions performed by these professionals and is not explicitly covered by another barrier assigned to the same initiating event. | Normal | TRUE | SIN
775 | 79 | BA-SIN-001 | Positioning and immobilization of the patient in the treatment position for special sessions (e.g., initial session), in the presence of the RTO and RTT, and whenever possible, the QMP. | Positioning and immobilization of the patient in the presence of the RTO and RTT, and whenever possible the QMP, following a procedure that defines the tasks of each professional involved. This step enables detection of geometric, dosimetric, or patient identification errors from previous stages of the treatment planning process. The barrier applies whenever the error can be identified through the functions performed by these professionals and is not explicitly covered by another barrier assigned to the same initiating event. | Normal | TRUE | SIN
776 | 81 | BA-SIN-001 | Positioning and immobilization of the patient in the treatment position for special sessions (e.g., initial session), in the presence of the RTO and RTT, and whenever possible, the QMP. | Positioning and immobilization of the patient in the presence of the RTO and RTT, and whenever possible the QMP, following a procedure that defines the tasks of each professional involved. This step enables detection of geometric, dosimetric, or patient identification errors from previous stages of the treatment planning process. The barrier applies whenever the error can be identified through the functions performed by these professionals and is not explicitly covered by another barrier assigned to the same initiating event. | Normal | TRUE | SIN
777 | 82 | BA-SIN-001 | Positioning and immobilization of the patient in the treatment position for special sessions (e.g., initial session), in the presence of the RTO and RTT, and whenever possible, the QMP. | Positioning and immobilization of the patient in the presence of the RTO and RTT, and whenever possible the QMP, following a procedure that defines the tasks of each professional involved. This step enables detection of geometric, dosimetric, or patient identification errors from previous stages of the treatment planning process. The barrier applies whenever the error can be identified through the functions performed by these professionals and is not explicitly covered by another barrier assigned to the same initiating event. | Normal | TRUE | SIN
778 | 66 | BA-SIN-002 | Customized positioning and immobilization devices. | Customized positioning and immobilization devices manufactured during anatomical data acquisition to ensure reproducibility and detect positioning errors. | Blanda | TRUE | SIN
779 | 9 | BA-SIN-003 | Verification of source-to-surface distance (SSD) and/or treatment table height. | Verification of correct positioning by checking the source-to-surface distance (SSD) and/or the isocenter height above the treatment table, enabling detection of positioning or geometric errors. | Blanda | TRUE | SIN
780 | 13 | BA-SIN-003 | Verification of source-to-surface distance (SSD) and/or treatment table height. | Verification of correct positioning by checking the source-to-surface distance (SSD) and/or the isocenter height above the treatment table, enabling detection of positioning or geometric errors. | Blanda | TRUE | SIN
781 | 14 | BA-SIN-003 | Verification of source-to-surface distance (SSD) and/or treatment table height. | Verification of correct positioning by checking the source-to-surface distance (SSD) and/or the isocenter height above the treatment table, enabling detection of positioning or geometric errors. | Blanda | TRUE | SIN
782 | 24 | BA-SIN-003 | Verification of source-to-surface distance (SSD) and/or treatment table height. | Verification of correct positioning by checking the source-to-surface distance (SSD) and/or the isocenter height above the treatment table, enabling detection of positioning or geometric errors. | Blanda | TRUE | SIN
783 | 26 | BA-SIN-003 | Verification of source-to-surface distance (SSD) and/or treatment table height. | Verification of correct positioning by checking the source-to-surface distance (SSD) and/or the isocenter height above the treatment table, enabling detection of positioning or geometric errors. | Blanda | TRUE | SIN
784 | 27 | BA-SIN-003 | Verification of source-to-surface distance (SSD) and/or treatment table height. | Verification of correct positioning by checking the source-to-surface distance (SSD) and/or the isocenter height above the treatment table, enabling detection of positioning or geometric errors. | Blanda | TRUE | SIN
785 | 38 | BA-SIN-003 | Verification of source-to-surface distance (SSD) and/or treatment table height. | Verification of correct positioning by checking the source-to-surface distance (SSD) and/or the isocenter height above the treatment table, enabling detection of positioning or geometric errors. | Blanda | TRUE | SIN
786 | 42 | BA-SIN-003 | Verification of source-to-surface distance (SSD) and/or treatment table height. | Verification of correct positioning by checking the source-to-surface distance (SSD) and/or the isocenter height above the treatment table, enabling detection of positioning or geometric errors. | Blanda | TRUE | SIN
787 | 64 | BA-SIN-003 | Verification of source-to-surface distance (SSD) and/or treatment table height. | Verification of correct positioning by checking the source-to-surface distance (SSD) and/or the isocenter height above the treatment table, enabling detection of positioning or geometric errors. | Blanda | TRUE | SIN
788 | 65 | BA-SIN-003 | Verification of source-to-surface distance (SSD) and/or treatment table height. | Verification of correct positioning by checking the source-to-surface distance (SSD) and/or the isocenter height above the treatment table, enabling detection of positioning or geometric errors. | Blanda | TRUE | SIN
789 | 66 | BA-SIN-003 | Verification of source-to-surface distance (SSD) and/or treatment table height. | Verification of correct positioning by checking the source-to-surface distance (SSD) and/or the isocenter height above the treatment table, enabling detection of positioning or geometric errors. | Blanda | TRUE | SIN
790 | 67 | BA-SIN-003 | Verification of source-to-surface distance (SSD) and/or treatment table height. | Verification of correct positioning by checking the source-to-surface distance (SSD) and/or the isocenter height above the treatment table, enabling detection of positioning or geometric errors. | Blanda | TRUE | SIN
791 | 68 | BA-SIN-003 | Verification of source-to-surface distance (SSD) and/or treatment table height. | Verification of correct positioning by checking the source-to-surface distance (SSD) and/or the isocenter height above the treatment table, enabling detection of positioning or geometric errors. | Blanda | TRUE | SIN
792 | 69 | BA-SIN-003 | Verification of source-to-surface distance (SSD) and/or treatment table height. | Verification of correct positioning by checking the source-to-surface distance (SSD) and/or the isocenter height above the treatment table, enabling detection of positioning or geometric errors. | Blanda | TRUE | SIN
793 | 70 | BA-SIN-003 | Verification of source-to-surface distance (SSD) and/or treatment table height. | Verification of correct positioning by checking the source-to-surface distance (SSD) and/or the isocenter height above the treatment table, enabling detection of positioning or geometric errors. | Blanda | TRUE | SIN
794 | 72 | BA-SIN-003 | Verification of source-to-surface distance (SSD) and/or treatment table height. | Verification of correct positioning by checking the source-to-surface distance (SSD) and/or the isocenter height above the treatment table, enabling detection of positioning or geometric errors. | Blanda | TRUE | SIN
795 | 83 | BA-SIN-003 | Verification of source-to-surface distance (SSD) and/or treatment table height. | Verification of correct positioning by checking the source-to-surface distance (SSD) and/or the isocenter height above the treatment table, enabling detection of positioning or geometric errors. | Blanda | TRUE | SIN
796 | 84 | BA-SIN-003 | Verification of source-to-surface distance (SSD) and/or treatment table height. | Verification of correct positioning by checking the source-to-surface distance (SSD) and/or the isocenter height above the treatment table, enabling detection of positioning or geometric errors. | Blanda | TRUE | SIN
797 | 85 | BA-SIN-003 | Verification of source-to-surface distance (SSD) and/or treatment table height. | Verification of correct positioning by checking the source-to-surface distance (SSD) and/or the isocenter height above the treatment table, enabling detection of positioning or geometric errors. | Blanda | TRUE | SIN
798 | 86 | BA-SIN-003 | Verification of source-to-surface distance (SSD) and/or treatment table height. | Verification of correct positioning by checking the source-to-surface distance (SSD) and/or the isocenter height above the treatment table, enabling detection of positioning or geometric errors. | Blanda | TRUE | SIN
799 | 87 | BA-SIN-003 | Verification of source-to-surface distance (SSD) and/or treatment table height. | Verification of correct positioning by checking the source-to-surface distance (SSD) and/or the isocenter height above the treatment table, enabling detection of positioning or geometric errors. | Blanda | TRUE | SIN
800 | 88 | BA-SIN-003 | Verification of source-to-surface distance (SSD) and/or treatment table height. | Verification of correct positioning by checking the source-to-surface distance (SSD) and/or the isocenter height above the treatment table, enabling detection of positioning or geometric errors. | Blanda | TRUE | SIN
801 | 89 | BA-SIN-003 | Verification of source-to-surface distance (SSD) and/or treatment table height. | Verification of correct positioning by checking the source-to-surface distance (SSD) and/or the isocenter height above the treatment table, enabling detection of positioning or geometric errors. | Blanda | TRUE | SIN
802 | 91 | BA-SIN-003 | Verification of source-to-surface distance (SSD) and/or treatment table height. | Verification of correct positioning by checking the source-to-surface distance (SSD) and/or the isocenter height above the treatment table, enabling detection of positioning or geometric errors. | Blanda | TRUE | SIN
803 | 9 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
804 | 10 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
805 | 11 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
806 | 12 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
807 | 13 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
808 | 14 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
809 | 15 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
810 | 16 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
811 | 18 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
812 | 19 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
813 | 24 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
814 | 26 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
815 | 27 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
816 | 28 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
817 | 29 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
818 | 30 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
819 | 31 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
820 | 32 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
821 | 33 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
822 | 34 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
823 | 37 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
824 | 38 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
825 | 41 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
826 | 42 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
827 | 53 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
828 | 56 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
829 | 64 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
830 | 65 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
831 | 66 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
832 | 67 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
833 | 68 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
834 | 69 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
835 | 70 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
836 | 71 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
837 | 72 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
838 | 80 | BA-SIN-004 | Analysis of verification images during the initial session and selected daily treatment sessions. | Analysis of verification images during the initial session and selected daily treatment sessions to detect potential geometric errors from previous stages. | Blanda | TRUE | SIN
839 | 3 | BA-SIN-005 | Preparation and review of information prior to the initial session by the RTT. | Preparation and review of all patient treatment information in the OIS prior to the initial session. The RTT ensures correct transfer of treatment fields, creation of the treatment schedule, and availability of verification images, among other checks. | Blanda | TRUE | SIN
840 | 6 | BA-SIN-005 | Preparation and review of information prior to the initial session by the RTT. | Preparation and review of all patient treatment information in the OIS prior to the initial session. The RTT ensures correct transfer of treatment fields, creation of the treatment schedule, and availability of verification images, among other checks. | Blanda | TRUE | SIN
841 | 39 | BA-SIN-005 | Preparation and review of information prior to the initial session by the RTT. | Preparation and review of all patient treatment information in the OIS prior to the initial session. The RTT ensures correct transfer of treatment fields, creation of the treatment schedule, and availability of verification images, among other checks. | Blanda | TRUE | SIN
842 | 50 | BA-SIN-005 | Preparation and review of information prior to the initial session by the RTT. | Preparation and review of all patient treatment information in the OIS prior to the initial session. The RTT ensures correct transfer of treatment fields, creation of the treatment schedule, and availability of verification images, among other checks. | Blanda | TRUE | SIN
843 | 83 | BA-SIN-006 | Daily imaging and analysis for IGRT. | Daily imaging and analysis for IGRT to detect patient identification errors, positioning inaccuracies, or anatomical changes. | Normal | TRUE | SIN
844 | 84 | BA-SIN-006 | Daily imaging and analysis for IGRT. | Daily imaging and analysis for IGRT to detect patient identification errors, positioning inaccuracies, or anatomical changes. | Normal | TRUE | SIN
845 | 85 | BA-SIN-006 | Daily imaging and analysis for IGRT. | Daily imaging and analysis for IGRT to detect patient identification errors, positioning inaccuracies, or anatomical changes. | Normal | TRUE | SIN
846 | 86 | BA-SIN-006 | Daily imaging and analysis for IGRT. | Daily imaging and analysis for IGRT to detect patient identification errors, positioning inaccuracies, or anatomical changes. | Normal | TRUE | SIN
847 | 87 | BA-SIN-006 | Daily imaging and analysis for IGRT. | Daily imaging and analysis for IGRT to detect patient identification errors, positioning inaccuracies, or anatomical changes. | Normal | TRUE | SIN
848 | 88 | BA-SIN-006 | Daily imaging and analysis for IGRT. | Daily imaging and analysis for IGRT to detect patient identification errors, positioning inaccuracies, or anatomical changes. | Normal | TRUE | SIN
849 | 89 | BA-SIN-006 | Daily imaging and analysis for IGRT. | Daily imaging and analysis for IGRT to detect patient identification errors, positioning inaccuracies, or anatomical changes. | Normal | TRUE | SIN
850 | 91 | BA-SIN-006 | Daily imaging and analysis for IGRT. | Daily imaging and analysis for IGRT to detect patient identification errors, positioning inaccuracies, or anatomical changes. | Normal | TRUE | SIN
851 | 92 | BA-SIN-006 | Daily imaging and analysis for IGRT. | Daily imaging and analysis for IGRT to detect patient identification errors, positioning inaccuracies, or anatomical changes. | Normal | TRUE | SIN
852 | 93 | BA-SIN-006 | Daily imaging and analysis for IGRT. | Daily imaging and analysis for IGRT to detect patient identification errors, positioning inaccuracies, or anatomical changes. | Normal | TRUE | SIN
853 | 102 | BA-SIN-006 | Daily imaging and analysis for IGRT. | Daily imaging and analysis for IGRT to detect patient identification errors, positioning inaccuracies, or anatomical changes. | Normal | TRUE | SIN
854 | 21 | BA-SIN-007 | Verification of image fusion results. | When automatic fusion tools are used, the fusion result must be reviewed by a qualified professional. If fusion is performed manually by one person, it must be independently checked by another professional. | Normal | TRUE | SIN
855 | 22 | BA-SIN-007 | Verification of image fusion results. | When automatic fusion tools are used, the fusion result must be reviewed by a qualified professional. If fusion is performed manually by one person, it must be independently checked by another professional. | Normal | TRUE | SIN
856 | 23 | BA-SIN-007 | Verification of image fusion results. | When automatic fusion tools are used, the fusion result must be reviewed by a qualified professional. If fusion is performed manually by one person, it must be independently checked by another professional. | Normal | TRUE | SIN
857 | 30 | BA-SIN-007 | Verification of image fusion results. | When automatic fusion tools are used, the fusion result must be reviewed by a qualified professional. If fusion is performed manually by one person, it must be independently checked by another professional. | Normal | TRUE | SIN
858 | 31 | BA-SIN-007 | Verification of image fusion results. | When automatic fusion tools are used, the fusion result must be reviewed by a qualified professional. If fusion is performed manually by one person, it must be independently checked by another professional. | Normal | TRUE | SIN
859 | 76 | BA-SIN-007 | Verification of image fusion results. | When automatic fusion tools are used, the fusion result must be reviewed by a qualified professional. If fusion is performed manually by one person, it must be independently checked by another professional. | Normal | TRUE | SIN
860 | 97 | BA-SIN-007 | Verification of image fusion results. | When automatic fusion tools are used, the fusion result must be reviewed by a qualified professional. If fusion is performed manually by one person, it must be independently checked by another professional. | Normal | TRUE | SIN
861 | 76 | BA-SIN-008 | Acquire verification images for each isocenter after all displacements have been completed and immediately before treatment. | Acquire verification images for each isocenter after all movements have been completed and immediately before treatment. These images are intended to confirm that the applied displacements are correct, not to adjust positioning. | Blanda | TRUE | SIN
862 | 77 | BA-SIN-008 | Acquire verification images for each isocenter after all displacements have been completed and immediately before treatment. | Acquire verification images for each isocenter after all movements have been completed and immediately before treatment. These images are intended to confirm that the applied displacements are correct, not to adjust positioning. | Blanda | TRUE | SIN
863 | 78 | BA-SIN-008 | Acquire verification images for each isocenter after all displacements have been completed and immediately before treatment. | Acquire verification images for each isocenter after all movements have been completed and immediately before treatment. These images are intended to confirm that the applied displacements are correct, not to adjust positioning. | Blanda | TRUE | SIN
864 | 80 | BA-SIN-008 | Acquire verification images for each isocenter after all displacements have been completed and immediately before treatment. | Acquire verification images for each isocenter after all movements have been completed and immediately before treatment. These images are intended to confirm that the applied displacements are correct, not to adjust positioning. | Blanda | TRUE | SIN
865 | 97 | BA-SIN-008 | Acquire verification images for each isocenter after all displacements have been completed and immediately before treatment. | Acquire verification images for each isocenter after all movements have been completed and immediately before treatment. These images are intended to confirm that the applied displacements are correct, not to adjust positioning. | Blanda | TRUE | SIN
866 | 98 | BA-SIN-008 | Acquire verification images for each isocenter after all displacements have been completed and immediately before treatment. | Acquire verification images for each isocenter after all movements have been completed and immediately before treatment. These images are intended to confirm that the applied displacements are correct, not to adjust positioning. | Blanda | TRUE | SIN
867 | 99 | BA-SIN-008 | Acquire verification images for each isocenter after all displacements have been completed and immediately before treatment. | Acquire verification images for each isocenter after all movements have been completed and immediately before treatment. These images are intended to confirm that the applied displacements are correct, not to adjust positioning. | Blanda | TRUE | SIN
868 | 25 | BA-SIN-009 | Pre-treatment verification that the phase and amplitude of respiratory motion match those recorded during the simulation. | Pre-treatment verification that the phase and amplitude of respiratory motion match those recorded during simulation. This is performed by monitoring several breathing cycles prior to treatment and maintained throughout the session. | Normal | TRUE | SIN
869 | 61 | BA-SIN-009 | Pre-treatment verification that the phase and amplitude of respiratory motion match those recorded during the simulation. | Pre-treatment verification that the phase and amplitude of respiratory motion match those recorded during simulation. This is performed by monitoring several breathing cycles prior to treatment and maintained throughout the session. | Normal | TRUE | SIN
870 | 94 | BA-SIN-009 | Pre-treatment verification that the phase and amplitude of respiratory motion match those recorded during the simulation. | Pre-treatment verification that the phase and amplitude of respiratory motion match those recorded during simulation. This is performed by monitoring several breathing cycles prior to treatment and maintained throughout the session. | Normal | TRUE | SIN
871 | 101 | BA-SIN-009 | Pre-treatment verification that the phase and amplitude of respiratory motion match those recorded during the simulation. | Pre-treatment verification that the phase and amplitude of respiratory motion match those recorded during simulation. This is performed by monitoring several breathing cycles prior to treatment and maintained throughout the session. | Normal | TRUE | SIN
872 | 16 | BA-TTD-001 | Verification of patient positioning and immobilization during each daily treatment session. | Verification by an RTT of patient positioning and immobilization during daily treatment to detect geometric errors, such as incorrect positioning or deficiencies in immobilization devices. | Blanda | TRUE | TTD
873 | 87 | BA-TTD-001 | Verification of patient positioning and immobilization during each daily treatment session. | Verification by an RTT of patient positioning and immobilization during daily treatment to detect geometric errors, such as incorrect positioning or deficiencies in immobilization devices. | Blanda | TRUE | TTD
874 | 61 | BA-TTD-002 | Comparison of immobilization, indexing, and positioning parameters (including respiratory monitoring) with reference values documented on the simulation sheet. | Verification of immobilization and indexing parameters used for patient positioning by comparing them with reference values on the simulation sheet, enabling detection of positioning errors. | Blanda | TRUE | TTD
875 | 66 | BA-TTD-002 | Comparison of immobilization, indexing, and positioning parameters (including respiratory monitoring) with reference values documented on the simulation sheet. | Verification of immobilization and indexing parameters used for patient positioning by comparing them with reference values on the simulation sheet, enabling detection of positioning errors. | Blanda | TRUE | TTD
876 | 67 | BA-TTD-002 | Comparison of immobilization, indexing, and positioning parameters (including respiratory monitoring) with reference values documented on the simulation sheet. | Verification of immobilization and indexing parameters used for patient positioning by comparing them with reference values on the simulation sheet, enabling detection of positioning errors. | Blanda | TRUE | TTD
877 | 69 | BA-TTD-002 | Comparison of immobilization, indexing, and positioning parameters (including respiratory monitoring) with reference values documented on the simulation sheet. | Verification of immobilization and indexing parameters used for patient positioning by comparing them with reference values on the simulation sheet, enabling detection of positioning errors. | Blanda | TRUE | TTD
878 | 83 | BA-TTD-002 | Comparison of immobilization, indexing, and positioning parameters (including respiratory monitoring) with reference values documented on the simulation sheet. | Verification of immobilization and indexing parameters used for patient positioning by comparing them with reference values on the simulation sheet, enabling detection of positioning errors. | Blanda | TRUE | TTD
879 | 85 | BA-TTD-002 | Comparison of immobilization, indexing, and positioning parameters (including respiratory monitoring) with reference values documented on the simulation sheet. | Verification of immobilization and indexing parameters used for patient positioning by comparing them with reference values on the simulation sheet, enabling detection of positioning errors. | Blanda | TRUE | TTD
880 | 86 | BA-TTD-002 | Comparison of immobilization, indexing, and positioning parameters (including respiratory monitoring) with reference values documented on the simulation sheet. | Verification of immobilization and indexing parameters used for patient positioning by comparing them with reference values on the simulation sheet, enabling detection of positioning errors. | Blanda | TRUE | TTD
881 | 88 | BA-TTD-002 | Comparison of immobilization, indexing, and positioning parameters (including respiratory monitoring) with reference values documented on the simulation sheet. | Verification of immobilization and indexing parameters used for patient positioning by comparing them with reference values on the simulation sheet, enabling detection of positioning errors. | Blanda | TRUE | TTD
882 | 93 | BA-TTD-002 | Comparison of immobilization, indexing, and positioning parameters (including respiratory monitoring) with reference values documented on the simulation sheet. | Verification of immobilization and indexing parameters used for patient positioning by comparing them with reference values on the simulation sheet, enabling detection of positioning errors. | Blanda | TRUE | TTD
883 | 94 | BA-TTD-002 | Comparison of immobilization, indexing, and positioning parameters (including respiratory monitoring) with reference values documented on the simulation sheet. | Verification of immobilization and indexing parameters used for patient positioning by comparing them with reference values on the simulation sheet, enabling detection of positioning errors. | Blanda | TRUE | TTD
884 | 101 | BA-TTD-002 | Comparison of immobilization, indexing, and positioning parameters (including respiratory monitoring) with reference values documented on the simulation sheet. | Verification of immobilization and indexing parameters used for patient positioning by comparing them with reference values on the simulation sheet, enabling detection of positioning errors. | Blanda | TRUE | TTD
885 | 58 | BA-TTD-003 | Daily verification of the cumulative dose at the reference dose point. | Verification of the accumulated dose at the dose point at the end of the session. | Blanda | TRUE | TTD
886 | 62 | BA-TTD-003 | Daily verification of the cumulative dose at the reference dose point. | Verification of the accumulated dose at the dose point at the end of the session. | Blanda | TRUE | TTD
887 | 104 | BA-TTD-003 | Daily verification of the cumulative dose at the reference dose point. | Verification of the accumulated dose at the dose point at the end of the session. | Blanda | TRUE | TTD
888 | 105 | BA-TTD-003 | Daily verification of the cumulative dose at the reference dose point. | Verification of the accumulated dose at the dose point at the end of the session. | Blanda | TRUE | TTD
889 | 106 | BA-TTD-003 | Daily verification of the cumulative dose at the reference dose point. | Verification of the accumulated dose at the dose point at the end of the session. | Blanda | TRUE | TTD
890 | 107 | BA-TTD-003 | Daily verification of the cumulative dose at the reference dose point. | Verification of the accumulated dose at the dose point at the end of the session. | Blanda | TRUE | TTD
891 | 53 | BA-TTD-004 | Software-based blocking of patient treatment as indicated by the RTO. | The radiation oncologist’s OIS block on the patient’s treatment prevents continuation until the modification is implemented and approved. | Robusta | TRUE | TTD
892 | 81 | BA-TTD-004 | Software-based blocking of patient treatment as indicated by the RTO. | The radiation oncologist’s OIS block on the patient’s treatment prevents continuation until the modification is implemented and approved. | Robusta | TRUE | TTD
893 | 43 | BA-VER-001 | Post-planning dosimetric verification, including verification of any subsequent modifications. | Measurement of treatment fields using various methods (e.g., ionization chambers, film, or diode arrays) enables detection and verification of errors in the accelerator, TPS, or planning parameters. Some systems allow these measurements to be back-projected onto the patient’s CT images, facilitating plan evaluation based on measured data. | Normal | TRUE | VER
894 | 49 | BA-VER-001 | Post-planning dosimetric verification, including verification of any subsequent modifications. | Measurement of treatment fields using various methods (e.g., ionization chambers, film, or diode arrays) enables detection and verification of errors in the accelerator, TPS, or planning parameters. Some systems allow these measurements to be back-projected onto the patient’s CT images, facilitating plan evaluation based on measured data. | Normal | TRUE | VER
895 | 52 | BA-VER-001 | Post-planning dosimetric verification, including verification of any subsequent modifications. | Measurement of treatment fields using various methods (e.g., ionization chambers, film, or diode arrays) enables detection and verification of errors in the accelerator, TPS, or planning parameters. Some systems allow these measurements to be back-projected onto the patient’s CT images, facilitating plan evaluation based on measured data. | Normal | TRUE | VER
896 | 56 | BA-VER-001 | Post-planning dosimetric verification, including verification of any subsequent modifications. | Measurement of treatment fields using various methods (e.g., ionization chambers, film, or diode arrays) enables detection and verification of errors in the accelerator, TPS, or planning parameters. Some systems allow these measurements to be back-projected onto the patient’s CT images, facilitating plan evaluation based on measured data. | Normal | TRUE | VER
897 | 59 | BA-VER-001 | Post-planning dosimetric verification, including verification of any subsequent modifications. | Measurement of treatment fields using various methods (e.g., ionization chambers, film, or diode arrays) enables detection and verification of errors in the accelerator, TPS, or planning parameters. Some systems allow these measurements to be back-projected onto the patient’s CT images, facilitating plan evaluation based on measured data. | Normal | TRUE | VER
898 | 82 | BA-VER-001 | Post-planning dosimetric verification, including verification of any subsequent modifications. | Measurement of treatment fields using various methods (e.g., ionization chambers, film, or diode arrays) enables detection and verification of errors in the accelerator, TPS, or planning parameters. Some systems allow these measurements to be back-projected onto the patient’s CT images, facilitating plan evaluation based on measured data. | Normal | TRUE | VER
899 | 62 | BA-VER-002 | System alert before irradiation when in verification mode. | The system provides an alert before irradiation in verification mode. | Robusta | TRUE | VER
900 | 63 | BA-VER-002 | System alert before irradiation when in verification mode. | The system provides an alert before irradiation in verification mode. | Robusta | TRUE | VER
901 | 58 | BA-VER-003 | User group permissions enabling operating in QA or verification mode. | Permissions that prevent RTTs from using their username and password in verification mode. | Muy robusta | TRUE | VER
902 | 62 | BA-VER-003 | User group permissions enabling operating in QA or verification mode. | Permissions that prevent RTTs from using their username and password in verification mode. | Muy robusta | TRUE | VER
903 | 63 | BA-VER-003 | User group permissions enabling operating in QA or verification mode. | Permissions that prevent RTTs from using their username and password in verification mode. | Muy robusta | TRUE | VER
904 | 58 | BA-VER-004 | Independent manual recording of daily treatment. | Independent manual recording of sessions, for example, using a card to track the number of treatments or patient check-ins. | Blanda | TRUE | VER
905 | 59 | BA-VER-005 | Joint replanning and evaluation of results by the RTO and QMP. | Joint replanning and evaluation of planning results by the RTO and QMP. | Normal | TRUE | VER
906 | 55 | BA-VER-006 | Treatment plan approval by the RTO following completion of dosimetric verification. | Treatment plan approval by the RTO after dosimetric verification, ensuring prevention of treatment with a non-verified plan. | Blanda | TRUE | VER
907 | 57 | BA-VER-006 | Treatment plan approval by the RTO following completion of dosimetric verification. | Treatment plan approval by the RTO after dosimetric verification, ensuring prevention of treatment with a non-verified plan. | Blanda | TRUE | VER
908 | 58 | BA-VER-006 | Treatment plan approval by the RTO following completion of dosimetric verification. | Treatment plan approval by the RTO after dosimetric verification, ensuring prevention of treatment with a non-verified plan. | Blanda | TRUE | VER

## Reductores de consecuencia
Id | Suceso | Código | Nombre | Descripción | Robustez | Marcado | Orden
909 | 71 | RC-CON-001 | Detection of systematic positioning deviations through image analysis, following established procedures. | Detection of systematic positioning deviations through image analysis, following established procedures. | Normal | TRUE | CON
910 | 91 | RC-CON-001 | Detection of systematic positioning deviations through image analysis, following established procedures. | Detection of systematic positioning deviations through image analysis, following established procedures. | Normal | TRUE | CON
911 | 92 | RC-CON-001 | Detection of systematic positioning deviations through image analysis, following established procedures. | Detection of systematic positioning deviations through image analysis, following established procedures. | Normal | TRUE | CON
912 | 18 | RC-CON-002 | QA tests of the image acquisition system | QA tests of the image acquisition system, including checks for displacement accuracy, image quality, and associated dose parameters. | Normal | TRUE | CON
913 | 74 | RC-CON-002 | QA tests of the image acquisition system | QA tests of the image acquisition system, including checks for displacement accuracy, image quality, and associated dose parameters. | Normal | TRUE | CON
914 | 49 | RC-CON-003 | Periodic quality control program of the TPS. | Periodic quality control program of the TPS. | Blanda | TRUE | CON
915 | 100 | RC-CON-004 | Periodic quality control checks that include verification of this parameter. | Periodic quality control checks that include verification of this parameter. | Blanda | TRUE | CON
916 | 70 | RC-CON-005 | Daily quality control checks that include verification of this parameter. | Daily quality control checks to detect equipment failures since the previous check. | Normal | TRUE | CON
917 | 71 | RC-CON-005 | Daily quality control checks that include verification of this parameter. | Daily quality control checks to detect equipment failures since the previous check. | Normal | TRUE | CON
918 | 74 | RC-CON-005 | Daily quality control checks that include verification of this parameter. | Daily quality control checks to detect equipment failures since the previous check. | Normal | TRUE | CON
919 | 78 | RC-CON-005 | Daily quality control checks that include verification of this parameter. | Daily quality control checks to detect equipment failures since the previous check. | Normal | TRUE | CON
920 | 91 | RC-CON-005 | Daily quality control checks that include verification of this parameter. | Daily quality control checks to detect equipment failures since the previous check. | Normal | TRUE | CON
921 | 92 | RC-CON-005 | Daily quality control checks that include verification of this parameter. | Daily quality control checks to detect equipment failures since the previous check. | Normal | TRUE | CON
922 | 95 | RC-CON-005 | Daily quality control checks that include verification of this parameter. | Daily quality control checks to detect equipment failures since the previous check. | Normal | TRUE | CON
923 | 96 | RC-CON-005 | Daily quality control checks that include verification of this parameter. | Daily quality control checks to detect equipment failures since the previous check. | Normal | TRUE | CON
924 | 99 | RC-CON-005 | Daily quality control checks that include verification of this parameter. | Daily quality control checks to detect equipment failures since the previous check. | Normal | TRUE | CON
925 | 100 | RC-CON-005 | Daily quality control checks that include verification of this parameter. | Daily quality control checks to detect equipment failures since the previous check. | Normal | TRUE | CON
926 | 16 | RC-CON-006 | Quality control of immobilization systems. | Quality control of immobilization systems. | Blanda | TRUE | CON
927 | 95 | RC-CON-007 | Periodic verification that the image isocenter aligns with the radiation isocenter. | Periodic verification that the image isocenter coincides with the radiation isocenter. | Blanda | TRUE | CON
928 | 1 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
929 | 2 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
930 | 3 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
931 | 4 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
932 | 5 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
933 | 6 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
934 | 7 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
935 | 8 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
936 | 9 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
937 | 10 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
938 | 11 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
939 | 13 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
940 | 14 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
941 | 15 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
942 | 16 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
943 | 24 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
944 | 26 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
945 | 27 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
946 | 28 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
947 | 29 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
948 | 30 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
949 | 31 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
950 | 32 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
951 | 33 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
952 | 34 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
953 | 35 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
954 | 36 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
955 | 37 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
956 | 38 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
957 | 39 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
958 | 40 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
959 | 41 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
960 | 42 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
961 | 46 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
962 | 47 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
963 | 48 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
964 | 49 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
965 | 50 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
966 | 51 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
967 | 52 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
968 | 53 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
969 | 54 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
970 | 55 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
971 | 56 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
972 | 57 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
973 | 59 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
974 | 62 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
975 | 63 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
976 | 64 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
977 | 65 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
978 | 66 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
979 | 67 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
980 | 68 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
981 | 69 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
982 | 71 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
983 | 72 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
984 | 73 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
985 | 76 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
986 | 78 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
987 | 79 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
988 | 80 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
989 | 81 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
990 | 82 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
991 | 100 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
992 | 103 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
993 | 104 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
994 | 105 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
995 | 107 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
996 | 108 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
997 | 109 | RC-DEV-001 | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Periodic medical and/or nursing review of the patient to detect errors in treatment administration or earlier stages. | Normal | TRUE | DEV
998 | 39 | RC-DEV-002 | Periodic review of the treatment chart. | Periodic review of the treatment chart, supported by checklists, scripts, or automated processes to ensure treatment is progressing according to plan. | Normal | TRUE | DEV
999 | 40 | RC-DEV-002 | Periodic review of the treatment chart. | Periodic review of the treatment chart, supported by checklists, scripts, or automated processes to ensure treatment is progressing according to plan. | Normal | TRUE | DEV
1000 | 50 | RC-DEV-002 | Periodic review of the treatment chart. | Periodic review of the treatment chart, supported by checklists, scripts, or automated processes to ensure treatment is progressing according to plan. | Normal | TRUE | DEV
1001 | 53 | RC-DEV-002 | Periodic review of the treatment chart. | Periodic review of the treatment chart, supported by checklists, scripts, or automated processes to ensure treatment is progressing according to plan. | Normal | TRUE | DEV
1002 | 54 | RC-DEV-002 | Periodic review of the treatment chart. | Periodic review of the treatment chart, supported by checklists, scripts, or automated processes to ensure treatment is progressing according to plan. | Normal | TRUE | DEV
1003 | 57 | RC-DEV-002 | Periodic review of the treatment chart. | Periodic review of the treatment chart, supported by checklists, scripts, or automated processes to ensure treatment is progressing according to plan. | Normal | TRUE | DEV
1004 | 62 | RC-DEV-002 | Periodic review of the treatment chart. | Periodic review of the treatment chart, supported by checklists, scripts, or automated processes to ensure treatment is progressing according to plan. | Normal | TRUE | DEV
1005 | 63 | RC-DEV-002 | Periodic review of the treatment chart. | Periodic review of the treatment chart, supported by checklists, scripts, or automated processes to ensure treatment is progressing according to plan. | Normal | TRUE | DEV
1006 | 65 | RC-DEV-002 | Periodic review of the treatment chart. | Periodic review of the treatment chart, supported by checklists, scripts, or automated processes to ensure treatment is progressing according to plan. | Normal | TRUE | DEV
1007 | 81 | RC-DEV-002 | Periodic review of the treatment chart. | Periodic review of the treatment chart, supported by checklists, scripts, or automated processes to ensure treatment is progressing according to plan. | Normal | TRUE | DEV
1008 | 106 | RC-DEV-002 | Periodic review of the treatment chart. | Periodic review of the treatment chart, supported by checklists, scripts, or automated processes to ensure treatment is progressing according to plan. | Normal | TRUE | DEV
1009 | 107 | RC-DEV-002 | Periodic review of the treatment chart. | Periodic review of the treatment chart, supported by checklists, scripts, or automated processes to ensure treatment is progressing according to plan. | Normal | TRUE | DEV
1010 | 108 | RC-DEV-002 | Periodic review of the treatment chart. | Periodic review of the treatment chart, supported by checklists, scripts, or automated processes to ensure treatment is progressing according to plan. | Normal | TRUE | DEV
1011 | 109 | RC-DEV-002 | Periodic review of the treatment chart. | Periodic review of the treatment chart, supported by checklists, scripts, or automated processes to ensure treatment is progressing according to plan. | Normal | TRUE | DEV
1012 | 9 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1013 | 10 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1014 | 11 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1015 | 12 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1016 | 13 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1017 | 14 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1018 | 15 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1019 | 16 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1020 | 18 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1021 | 19 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1022 | 24 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1023 | 26 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1024 | 27 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1025 | 28 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1026 | 29 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1027 | 30 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1028 | 31 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1029 | 32 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1030 | 33 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1031 | 34 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1032 | 38 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1033 | 41 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1034 | 42 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1035 | 53 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1036 | 56 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1037 | 64 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1038 | 65 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1039 | 66 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1040 | 67 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1041 | 68 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1042 | 69 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1043 | 70 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1044 | 71 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1045 | 72 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1046 | 76 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1047 | 77 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1048 | 78 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1049 | 80 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1050 | 83 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1051 | 84 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1052 | 85 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1053 | 86 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1054 | 87 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1055 | 88 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1056 | 89 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1057 | 91 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1058 | 92 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1059 | 93 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1060 | 97 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1061 | 98 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1062 | 99 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1063 | 102 | RC-DEV-003 | Offline review of the IGRT image. | Offline review of the IGRT image by the RTO. | Normal | TRUE | DEV
1064 | 110 | RC-DEV-004 | Maintain working conditions that promote concentration and minimize distractions (e.g., designated areas or times without interruptions). | Maintain working conditions that promote concentration and minimize distractions (e.g., designated areas or times without interruptions). | Blanda | TRUE | DEV
1065 | 80 | RC-TTD-001 | Verification of the final isocenter marking on the day after treatment begins, following established procedures. | Verification of the final isocenter marking on the day after treatment begins, following established procedures. | Blanda | TRUE | TTD
1066 | 2 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1067 | 42 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1068 | 43 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1069 | 44 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1070 | 49 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1071 | 53 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1072 | 54 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1073 | 55 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1074 | 56 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1075 | 63 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1076 | 64 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1077 | 65 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1078 | 66 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1079 | 67 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1080 | 68 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1081 | 69 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1082 | 70 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1083 | 71 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1084 | 72 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1085 | 73 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1086 | 81 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1087 | 82 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1088 | 83 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1089 | 84 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1090 | 85 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1091 | 86 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1092 | 87 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1093 | 88 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1094 | 89 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1095 | 90 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1096 | 91 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1097 | 92 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1098 | 93 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1099 | 100 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1100 | 102 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1101 | 103 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1102 | 104 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1103 | 105 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1104 | 106 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1105 | 109 | RC-TTD-002 | Evaluation of dosimetry using the CBCT images. | Evaluation of dosimetry using the CBCT images. | Normal | TRUE | TTD
1106 | 110 | RC-TTD-003 | Observation of the treatment process via closed-circuit TV and activation of the emergency button if necessary. | Observation of the treatment process via closed-circuit TV and activation of the emergency button if necessary. | Muy blanda | TRUE | TTD
1107 | 49 | RC-TTD-004 | Detection of geometry or dose errors through visual indicators (e.g., skin discoloration) during daily positioning checks. | Detection of geometry or dose errors through visual indicators (e.g., skin discoloration) during daily positioning checks. | Blanda | TRUE | TTD
1108 | 56 | RC-TTD-004 | Detection of geometry or dose errors through visual indicators (e.g., skin discoloration) during daily positioning checks. | Detection of geometry or dose errors through visual indicators (e.g., skin discoloration) during daily positioning checks. | Blanda | TRUE | TTD
1109 | 110 | RC-TTD-005 | Use of the intercom in the control room to enable two-way communication between the patient and the RTT for treatment interruption if necessary. | Use of the intercom in the control room to enable two-way communication between the patient and the RTT for treatment interruption if necessary. | Muy blanda | TRUE | TTD
1110 | 39 | RC-TTD-006 | Daily consultation or printing of the session sheet in accordance with established procedures. | Daily consultation or printing of the session sheet in accordance with established procedures. | Blanda | TRUE | TTD
1111 | 39 | RC-TTD-007 | Independent manual recording of daily treatment by the RTT, following established procedures. | Independent manual recording of daily treatment by the RTT, following established procedures. | Blanda | TRUE | TTD
1112 | 107 | RC-TTD-007 | Independent manual recording of daily treatment by the RTT, following established procedures. | Independent manual recording of daily treatment by the RTT, following established procedures. | Blanda | TRUE | TTD
1113 | 108 | RC-TTD-007 | Independent manual recording of daily treatment by the RTT, following established procedures. | Independent manual recording of daily treatment by the RTT, following established procedures. | Blanda | TRUE | TTD
1114 | 109 | RC-TTD-007 | Independent manual recording of daily treatment by the RTT, following established procedures. | Independent manual recording of daily treatment by the RTT, following established procedures. | Blanda | TRUE | TTD
1115 | 44 | RC-TTD-008 | Treatment monitoring of beam flucence. | System that monitors beam fluence in real time (above or below the patient) and alerts to deviations from the planned dose. | Robusta | TRUE | TTD
1116 | 63 | RC-TTD-008 | Treatment monitoring of beam flucence. | System that monitors beam fluence in real time (above or below the patient) and alerts to deviations from the planned dose. | Robusta | TRUE | TTD
1117 | 100 | RC-TTD-008 | Treatment monitoring of beam flucence. | System that monitors beam fluence in real time (above or below the patient) and alerts to deviations from the planned dose. | Robusta | TRUE | TTD
1118 | 105 | RC-TTD-008 | Treatment monitoring of beam flucence. | System that monitors beam fluence in real time (above or below the patient) and alerts to deviations from the planned dose. | Robusta | TRUE | TTD
1119 | 106 | RC-TTD-008 | Treatment monitoring of beam flucence. | System that monitors beam fluence in real time (above or below the patient) and alerts to deviations from the planned dose. | Robusta | TRUE | TTD
1120 | 109 | RC-TTD-008 | Treatment monitoring of beam flucence. | System that monitors beam fluence in real time (above or below the patient) and alerts to deviations from the planned dose. | Robusta | TRUE | TTD
1121 | 25 | RC-TTD-009 | Periodic review of respiratory treatment records. | Periodic review of respiratory treatment records and comparison with reference values. | Normal | TRUE | TTD
1122 | 61 | RC-TTD-009 | Periodic review of respiratory treatment records. | Periodic review of respiratory treatment records and comparison with reference values. | Normal | TRUE | TTD
1123 | 94 | RC-TTD-009 | Periodic review of respiratory treatment records. | Periodic review of respiratory treatment records and comparison with reference values. | Normal | TRUE | TTD
1124 | 100 | RC-TTD-009 | Periodic review of respiratory treatment records. | Periodic review of respiratory treatment records and comparison with reference values. | Normal | TRUE | TTD